## Sheet1
# | Test Category | Component/Sub-System/System Involved | Test Name | Test Case Description | Test Set-Up | Acceptance Criteria | Automated | Results | Test Pass/Fail | Test Status | Test Owner | Planned Execution Date
1 | HW–Mechanical Integration | On/Off switch + Enclosure | Switch fit, travel, and sealing verification | 1) Assemble switch into enclosure cut-out. 2) Visually verify switch alignment and protrusion. 3) Press/release repeatedly and confirm free movement (no sticking). 4) Inspect PCB/wiring for stress while actuating the switch. | Fully assembled enclosure with switch installed | Switch is properly aligned, moves freely, protrudes correctly, sealing remains intact, and actuation does not stress PCB/wiring. |  |  |  |  | Abhilash
2 | HW–Electrical Integration | On/Off switch + Power-latch circuit | Debounced switch waveform and press-duration differentiation | 1) Power device from mains; connect oscilloscope to switch/power-latch signals. 2) Perform short press; capture waveform. 3) Perform long press; capture waveform. 4) Repeat on battery power. 5) Review waveforms for debounced behavior and distinct short vs long press behavior. | DUT; mains + battery operation; oscilloscope | Debounced signals are observed and short vs long presses are differentiable on both mains and battery, as per design intent/specification. |  |  |  |  | Abhilash
3 | HW–FW Integration | Power-latch + Firmware shutdown sequencing | Power-latch hold until safe shutdown completes | 1) Boot device on mains. 2) Initiate short-press shutdown. 3) Observe that power remains latched while shutdown proceeds. 4) Confirm power is released only after safe shutdown completion. 5) Repeat on battery and during mains↔battery transition. | DUT; mains + battery; ability to transition power sources; logging/indication available | Power-latch holds power until firmware completes safe shutdown; behavior remains correct on mains, on battery, and during transitions. |  |  |  |  | Abhilash
4 | HW–Mechanical Integration | 2.5 mm barrel connector + Enclosure | Barrel connector insertion and retention | 1) Insert barrel connector through enclosure opening. 2) Verify it can be secured/retained as designed. 3) Repeat insertion/removal cycles. 4) Inspect for loosening or damage. | DUT enclosure with power input hardware | Connector fits through opening and is retained with mechanical stability as designed; no damage after repeated use. |  |  |  |  | Abhilash
5 | Power-Path Integration | AC adapter input + Battery + PMIC + Firmware | Seamless switchover: mains → battery | 1) Run device on mains in normal operation. 2) Unplug mains while running. 3) Observe device continuity (no reset) and capture PMIC telemetry/logs. 4) Repeat multiple times. | DUT; controllable mains plug/unplug; logging/telemetry | Device remains operable without reboot; switchover occurs seamlessly and stays within design limits. |  |  |  |  | Abhilash
6 | Power-Path Integration | AC adapter input + Battery + PMIC + Firmware | Seamless switchover: battery → mains + charging state entry | 1) Run device on battery. 2) Plug in mains connector through enclosure. 3) Observe electrical transition to charge mode and capture PMIC telemetry/logs. 4) Repeat plug/unplug cycles. | DUT; battery operation; controllable mains; logging/telemetry | Mechanical insertion is smooth; system transitions to charging behavior without glitches; repeated plug/unplug produces no connector/power issues. |  |  |  |  | Abhilash
Automated Integration Testing
7 | HW–FW Integration | MAX30009 + firmware service + RC model circuit + analysis | MAX30009 Cole/Cole BCA automated integration | 1) Clear the base table if previously generated. 2) Connect MAX30009 module to the test RC model circuit. 3) Generate the base table and verify the “build base table started” response. 4) Configure sensor: fs = 400 Hz, current = 64 µA, stimulation frequency loops through 500, 1k, 2k, 5k, 10k, 20k, 50k, 100k, 200k, 300k, 450 kHz (5 s each). 5) For each frequency change, verify the firmware reports the measurement state sequence. 6) Request data every 0.5 s and record for 5 s at each frequency. 7) Record time-to-complete for each frequency segment. 8) Plot R vs f, Z vs f, Xc vs f, PhA vs f and compare mean error per frequency. 9) Plot Cole-Cole, extract R0 and Rinf, and calculate mean error for R0 and Rinf. 10) Power off MAX30009 and verify power-off response. | DUT with MAX30009 service available; RC model circuit; data capture + plotting environment | “build_base_table_started” is received; for every spectroscopy frequency change, firmware reports: pre_measuring → pre_measure_end (with real) → calibrating → calibrate_end → start_measuring; data is recorded for each frequency with time-to-complete noted; required plots are generated and mean errors are calculated and reported; MAX30009 power-off completes and is confirmed by the power-off response. Impedance ≤ 3 Ω; Resistance MAE ≤ 2 Ω; Reactance MAE ≤ 1 Ω; Phase angle MAE ≤ 0.2°. |  |  |  |  | Ansh and Kranti
8 | HW–FW Integration | MAX30009 + firmware service + resistor loads | MAX30009 ICG automated integration | 1) Clear base table if previously generated. 2) Generate the base table and confirm “build base table started”. 3) Connect MAX30009 ICG electrodes to a test resistor load (20Ω, 100Ω, 200Ω, 300Ω). 4) Configure: fs = 400 Hz, current = 1.28 µA, stimulation frequency = 20 kHz. 5) Verify the firmware reports the required measurement state sequence. 6) Request data every 0.5 s and record for 60 s. 7) Note the time taken to complete recording. 8) Calculate mean error for R and Z. 9) Extract synchronization markers and verify monotonic increase every 1 s. 10) Verify mean sampling frequency after 1 min is 400 Hz with error of 1 Hz. 11) Repeat across the other loads. 12) Power off MAX30009 and verify power-off response. | DUT with MAX30009 service available; selectable resistor loads; data capture/analysis | “build_base_table_started” is received; firmware reports the state sequence: pre_measuring → pre_measure_end (with real) → calibrating → calibrate_end → start_measuring; 60 s recording completes; mean error for R and Z is calculated and reported; synchronization markers are monotonically increasing every 1 s; mean sampling frequency after 1 min is 400 Hz with error of 1 Hz; MAX30009 power-off is confirmed by the power-off response. Impedance ≤ 3 Ω; Resistance MAE ≤ 2 Ω; Reactance MAE ≤ 1 Ω; Phase angle MAE ≤ 0.2°. |  |  |  |  | Ansh and Kranti
9 | HW–FW Integration | MAX30009 + firmware service + resistor load | MAX30009 ICG long-duration automated integration | 1) Clear base table if previously generated. 2) Generate base table and confirm “build base table started”. 3) Connect MAX30009 ICG electrodes to 100Ω test resistor. 4) Configure: fs = 400 Hz, current = 1.28 µA, stimulation frequency = 20 kHz. 5) Verify the firmware reports the required measurement state sequence. 6) Request data every 0.5 s and record for 1 hour. 7) Note time-to-complete recording. 8) Calculate mean error for R and Z. 9) Verify synchronization markers monotonically increase every 1 s. 10) Verify mean sampling frequency after 1 hour is 400 Hz with error of 1 Hz. 11) Power off MAX30009 and verify power-off response. | DUT with MAX30009 service available; 100Ω load; long-duration capture/analysis environment | “build_base_table_started” is received; firmware reports the measurement state sequence; continuous 1-hour recording completes; mean error for R and Z is calculated and reported; synchronization markers are monotonically increasing every 1 s; mean sampling frequency after 1 hour is 400 Hz with error of 1 Hz; MAX30009 power-off is confirmed by the power-off response. Impedance ≤ 3 Ω; Resistance MAE ≤ 2 Ω; Reactance MAE ≤ 1 Ω; Phase angle MAE ≤ 0.2°. |  |  |  |  | Ansh and Kranti
10 | HW–FW Integration | ADS1293 + firmware service + ECG simulator | ADS1293 ECG automated integration (60 s) | 1) Connect ADS1293 ECG inputs to an ECG simulator and set BPM = 30, 60, 120, 180 (as applicable to runs). 2) Configure ADS1293 such that ECG sampling frequency is 400 Hz and enable conversion. 3) Verify firmware returns “actual settings” and conversion is enabled. 4) Request data every 0.5 s and record continuously for 60 s. 5) Extract sync counter values and verify they are monotonically increasing every 1 s with no missing sync frames during the recording. 6) Calculate effective sampling frequency from total samples over 60 s and verify mean sampling frequency is 400 Hz with error not exceeding ±1 Hz. 7) Verify ECG signal integrity by matching BPM. 8) Log total sample count, timestamps, sync counter values, calculated sampling frequency, and signal quality metrics. 9) Stop conversion and power off ADS1293; verify power-off response. | DUT with ADS1293 service available; ECG simulator; capture/logging/reporting environment | Firmware returns “actual settings” with conversion enabled; 60 s recording completes; sync counters monotonically increase every 1 s with no missing sync frames; mean sampling frequency is 400 Hz with error not exceeding ±1 Hz; BPM matches simulator; required metrics are logged; ADS1293 power-off is confirmed by the power-off response. The device shall record ECG with the following accuracy metrics a. ±25uV for reference values less than or equal to 500 uV b. 5% or ±40uV for reference values greater than 500uV BPM Error = ±3 bpm or ±1%, whichever is greater |  |  |  |  | Ansh and Kranti
11 | HW–FW Integration | ADS1293 + firmware service + ECG simulator | ADS1293 ECG long-duration automated integration (1 hr) | 1) Connect ECG simulator; set BPM = 30, 60, 120, 180 (as applicable to runs). 2) Configure ADS1293 such that ECG sampling frequency is 400 Hz and enable conversion. 3) Verify firmware returns “actual settings” and conversion is enabled. 4) Request data every 0.5 s and record continuously for 1 hour. 5) Verify sync counters monotonically increase every 1 s with no missing sync frames during the recording. 6) Calculate effective sampling frequency from total samples over 1 hour and verify mean sampling frequency is 400 Hz with error not exceeding ±1 Hz. 7) Verify ECG signal integrity by matching BPM. 8) Log sample count, timestamps, sync counters, sampling frequency, and signal quality metrics. 9) Stop conversion and power off ADS1293; verify power-off response. | DUT with ADS1293 service available; ECG simulator; long-duration capture/logging | Firmware returns “actual settings” with conversion enabled; 60 s recording completes; sync counters monotonically increase every 1 s with no missing sync frames; mean sampling frequency is 400 Hz with error not exceeding ±1 Hz; BPM matches simulator; required metrics are logged; ADS1293 power-off is confirmed by the power-off response. The device shall record ECG with the following accuracy metrics a. ±25uV for reference values less than or equal to 500 uV b. 5% or ±40uV for reference values greater than 500uV BPM Error = ±3 bpm or ±1%, whichever is greater |  |  |  |  | Ansh and Kranti
12 | HW–FW Integration | ADS1293 + MAX30009 + firmware services + ECG simulator + ICG load | ADS1293 + MAX30009 synchronized acquisition (AC and Battery, 1 hr) | 1) Power system using AC adapter and confirm both services are reachable. 2) Connect ADS1293 inputs to ECG simulator (60 BPM). 3) Connect MAX30009 ICG electrodes to 100Ω resistor. 4) Configure ADS1293 for 400 Hz and enable conversion; confirm “actual settings”. 5) Configure MAX30009 for fs = 400 Hz, stimulation current = 1.28 mA, stimulation frequency = 20 kHz; confirm required meas_state sequence through start_measuring. 6) Start synchronized acquisition by requesting data from both services every 0.5 s. 7) Record ECG and ICG continuously for 1 hour. 8) Extract sync magic number and sync counter values from both streams and verify monotonic increase every 1 s with no missing/duplicated sync frames. 9) Calculate effective sampling frequency for both devices over 1 hour and verify mean FS = 400 Hz ± 1 Hz for each. 10) Compute clock drift between ADS1293 and MAX30009 using sync counter alignment and report total drift after 1 hour. 11) Perform FFT on ECG data (AC mode) and compute power ratio of signal band (0.7–3 Hz) vs mains interference (50 Hz); log ratio. 12) Verify ECG BPM matches simulator BPM. 13) Power down both devices and confirm power-off responses. 14) Switch system to battery mode. 15) Connect MAX30009 electrodes to patient and repeat the acquisition and analyses for a 1-hour recording on battery. 16) Perform FFT on ECG data (battery mode) and compute 0.7–3 Hz vs 50 Hz ratio. 17) Run ICG through SQA model and verify pass/fail per model output. 18) Log the final report items listed in the test definition. | DUT with both services; ECG simulator; resistor load; ability to operate on AC and battery; data processing environment | Both services reachable; ADS1293 “actual settings” returned; MAX30009 meas_state sequence observed through start_measuring; over 1 hour: sync counters monotonically increase every 1 s with no missing/duplicated sync frames; mean sampling frequency for both ADS1293 and MAX30009 is 400 Hz ± 1 Hz; clock drift is computed and reported; FFT-based 50 Hz noise power ratio is computed and logged (AC and battery); ECG BPM matches simulator; SQA model yields pass/fail output and is logged; all listed report metrics are captured; both devices confirm power-off responses. Error MAX30009: Impedance ≤ 3 Ω; Resistance MAE ≤ 2 Ω; Reactance MAE ≤ 1 Ω; Phase angle MAE ≤ 0.2°. Error ECG:The device shall record ECG with the following accuracy metrics a. ±25uV for reference values less than or equal to 500 uV b. 5% or ±40uV for reference values greater than 500uV BPM Error = ±3 bpm or ±1%, whichever is greater |  |  |  |  | Ansh and Kranti
13 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Integrity Baseline | 1) Create a test file of the using random data while computing a SHA256 checksum during write. 3) Write to storage. 4) Read the file back and compute SHA256 checksum again. 5) Compare expected vs read-back checksum. 6) Record PASS/FAIL result into the results log. | DUT (CM4) with eMMC mounted and writable; stress suite logging enabled (JSONL results). | PASS if read-back SHA256 checksum matches the write-time SHA256 checksum. FAIL if mismatch or if read/write errors occur. |  |  |  |  | Ansh and Kranti
14 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Endurance | 1) Start endurance mode for the 1-hr. 2) Repeatedly overwrite a single large file of the defined size until duration completes. 3) Ensure writes are flushed/synced at the end of each cycle. 4) Record the number of completed cycles. 5) Record PASS/FAIL into results log. | DUT (CM4) with sufficient free space on eMMC; stable power; logging enabled. | PASS if the endurance run completes for the full duration without OSError/exception. FAIL if storage I/O errors occur during the run. |  |  |  |  | Ansh and Kranti
15 | HW–FW Integration | CM4 eMMC + OS logging I/O + acquisition simulator thread | eMMC I/O Latency Under Logging | 1) Start the I/O stress test. 2) Start the background acquisition simulator that generates ECG (3×24-bit) + ICG (4×24-bit) samples at 400 Hz and flushes to disk periodically. 3) In parallel, generate repeated bursts of log append activity to storage with forced flush/sync. 4) Stop the simulator at end of duration. 5) Evaluate simulator timing metrics: achieved sample rate and max lag vs schedule. | DUT (CM4) with eMMC writable; stress suite running; logging enabled; no requirement for external equipment. | PASS if (a) achieved sample rate error is ≤ 1Hz relative to expected 400 Hz |  |  |  |  | Ansh and Kranti
16 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Integrity Post-Stress | 1) After endurance and I/O stress complete, run the same checksum-based integrity procedure again on a new test file. 2) Write random data while computing SHA256, sync writes, read back, recompute SHA256, compare. 3) Record PASS/FAIL into results log. | DUT (CM4) with eMMC mounted; stress suite logging enabled. | PASS if post-stress read-back SHA256 matches the write-time SHA256 checksum. FAIL if mismatch or if read/write errors occur. |  |  |  |  | Ansh and Kranti
17 | HW–FW Integration | CM4 SDRAM + OS + memtester utility (if installed) | SDRAM Memtester | 1) Launch the SDRAM memtester test. 2) Run memtester using the configured memory size and loop count. 3) Capture return code and record PASS/FAIL into results log. | DUT (CM4); memtester installed and executable; logging enabled. | PASS if memtester returns exit code 0 and reports no errors. FAIL if memtester missing or returns non-zero exit code. |  |  |  |  | Ansh and Kranti
18 | HW–FW Integration | CM4 SDRAM + OS memory manager + stress-ng (if available) or Python allocator | SDRAM Memory Pressure | 1) Start memory pressure test for the defined duration. 2) If stress-ng is available, run memory pressure targeting ~90% memory for the test duration; otherwise allocate memory in Python in 1 MB chunks until duration completes (or MemoryError occurs). 3) Release allocated memory at end. 4) Record PASS/FAIL into results log. | DUT (CM4); stress-ng optional; logging enabled. | PASS if the pressure run completes to duration (stress-ng exit code 0), or (fallback) completes and a MemoryError is treated as acceptable evidence of high pressure. FAIL if the stress tool exits non-zero or other exceptions occur. |  |  |  |  | Ansh and Kranti
19 | HW–FW Integration | CM4 CPU + thermal sensor + OS telemetry + stress-ng (if available) | CPU Thermal Load (HW-CPU-001-A) | 1) Start CPU thermal test for the defined duration. 2) Apply sustained CPU load (stress-ng if present; otherwise CPU-burn threads). 3) At the defined log interval, record CPU temperature and CPU usage telemetry to the log. 4) Run until duration completes. 5) Record PASS/FAIL into results log. | DUT (CM4); logging enabled; stable environment for duration. | PASS if the test runs for full duration and logs periodic temperature/CPU usage without exception. (Throttling/overheat determination is made by reviewing the logs produced by the script.) |  |  |  |  | Ansh and Kranti
20 | HW–FW Integration | ECG/ICG acquisition simulator + CM4 CPU load + eMMC logging I/O + OS scheduler | Full Concurrency (HW-SDRAM-001-D / HW-CPU-001-B) | 1) Start full concurrency test for the defined duration. 2) Run the ECG/ICG simulator at 400 Hz writing binary stream files. 3) Run CPU burn concurrently. 4) Generate continuous disk logging bursts with forced flush/sync while also recording temperature/CPU telemetry. 5) Stop simulator and CPU load after duration. 6) Evaluate simulator rate and max lag vs schedule. 7) Record PASS/FAIL with metrics into results log. | DUT (CM4) with eMMC writable; logging enabled. | PASS if simulator rate error ≤ 1Hz relative to 400Hz under combined CPU + disk I/O load. FAIL otherwise or on exceptions. |  |  |  |  | Ansh and Kranti
21 | HW–FW Integration | ADS1293 + firmware service + ECG simulator | ADS1293 ECG long-duration automated integration (1 hr) | 1) Connect ECG simulator; set BPM = 30, 60, 120, 180 (as applicable to runs). 2) Configure ADS1293 such that ECG sampling frequency is 400 Hz and enable conversion. 3) Verify firmware returns “actual settings” and conversion is enabled. 4) Request data every 0.5 s and record continuously for 1 hour. 5) Verify sync counters monotonically increase every 1 s with no missing sync frames during the recording. 6) Calculate effective sampling frequency from total samples over 1 hour and verify mean sampling frequency is 400 Hz with error not exceeding ±1 Hz. 7) Verify ECG signal integrity by matching BPM. 8) Log sample count, timestamps, sync counters, sampling frequency, and signal quality metrics. 9) Stop conversion and power off ADS1293; verify power-off response. | DUT with ADS1293 service available; ECG simulator; long-duration capture/logging | Firmware returns “actual settings” with conversion enabled; 60 s recording completes; sync counters monotonically increase every 1 s with no missing sync frames; mean sampling frequency is 400 Hz with error not exceeding ±1 Hz; BPM matches simulator; required metrics are logged; ADS1293 power-off is confirmed by the power-off response. The device shall record ECG with the following accuracy metrics a. ±25uV for reference values less than or equal to 500 uV b. 5% or ±40uV for reference values greater than 500uV BPM Error = ±3 bpm or ±1%, whichever is greater |  |  |  |  | Ansh and Kranti
22 | HW–FW Integration | ADS1293 + MAX30009 + firmware services + ECG simulator + ICG load | ADS1293 + MAX30009 synchronized acquisition (AC and Battery, 1 hr) | 1) Power system using AC adapter and confirm both services are reachable. 2) Connect ADS1293 inputs to ECG simulator (60 BPM). 3) Connect MAX30009 ICG electrodes to 100Ω resistor. 4) Configure ADS1293 for 400 Hz and enable conversion; confirm “actual settings”. 5) Configure MAX30009 for fs = 400 Hz, stimulation current = 1.28 mA, stimulation frequency = 20 kHz; confirm required meas_state sequence through start_measuring. 6) Start synchronized acquisition by requesting data from both services every 0.5 s. 7) Record ECG and ICG continuously for 1 hour. 8) Extract sync magic number and sync counter values from both streams and verify monotonic increase every 1 s with no missing/duplicated sync frames. 9) Calculate effective sampling frequency for both devices over 1 hour and verify mean FS = 400 Hz ± 1 Hz for each. 10) Compute clock drift between ADS1293 and MAX30009 using sync counter alignment and report total drift after 1 hour. 11) Perform FFT on ECG data (AC mode) and compute power ratio of signal band (0.7–3 Hz) vs mains interference (50 Hz); log ratio. 12) Verify ECG BPM matches simulator BPM. 13) Power down both devices and confirm power-off responses. 14) Switch system to battery mode. 15) Connect MAX30009 electrodes to patient and repeat the acquisition and analyses for a 1-hour recording on battery. 16) Perform FFT on ECG data (battery mode) and compute 0.7–3 Hz vs 50 Hz ratio. 17) Run ICG through SQA model and verify pass/fail per model output. 18) Log the final report items listed in the test definition. | DUT with both services; ECG simulator; resistor load; ability to operate on AC and battery; data processing environment | Both services reachable; ADS1293 “actual settings” returned; MAX30009 meas_state sequence observed through start_measuring; over 1 hour: sync counters monotonically increase every 1 s with no missing/duplicated sync frames; mean sampling frequency for both ADS1293 and MAX30009 is 400 Hz ± 1 Hz; clock drift is computed and reported; FFT-based 50 Hz noise power ratio is computed and logged (AC and battery); ECG BPM matches simulator; SQA model yields pass/fail output and is logged; all listed report metrics are captured; both devices confirm power-off responses. Error MAX30009: Impedance ≤ 3 Ω; Resistance MAE ≤ 2 Ω; Reactance MAE ≤ 1 Ω; Phase angle MAE ≤ 0.2°. Error ECG:The device shall record ECG with the following accuracy metrics a. ±25uV for reference values less than or equal to 500 uV b. 5% or ±40uV for reference values greater than 500uV BPM Error = ±3 bpm or ±1%, whichever is greater |  |  |  |  | Ansh and Kranti
23 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Integrity Baseline | 1) Create a test file of the using random data while computing a SHA256 checksum during write. 3) Write to storage. 4) Read the file back and compute SHA256 checksum again. 5) Compare expected vs read-back checksum. 6) Record PASS/FAIL result into the results log. | DUT (CM4) with eMMC mounted and writable; stress suite logging enabled (JSONL results). | PASS if read-back SHA256 checksum matches the write-time SHA256 checksum. FAIL if mismatch or if read/write errors occur. |  |  |  |  | Ansh and Kranti
24 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Endurance | 1) Start endurance mode for the 1-hr. 2) Repeatedly overwrite a single large file of the defined size until duration completes. 3) Ensure writes are flushed/synced at the end of each cycle. 4) Record the number of completed cycles. 5) Record PASS/FAIL into results log. | DUT (CM4) with sufficient free space on eMMC; stable power; logging enabled. | PASS if the endurance run completes for the full duration without OSError/exception. FAIL if storage I/O errors occur during the run. |  |  |  |  | Ansh and Kranti
25 | HW–FW Integration | CM4 eMMC + OS logging I/O + acquisition simulator thread | eMMC I/O Latency Under Logging | 1) Start the I/O stress test. 2) Start the background acquisition simulator that generates ECG (3×24-bit) + ICG (4×24-bit) samples at 400 Hz and flushes to disk periodically. 3) In parallel, generate repeated bursts of log append activity to storage with forced flush/sync. 4) Stop the simulator at end of duration. 5) Evaluate simulator timing metrics: achieved sample rate and max lag vs schedule. 6) Record PASS/FAIL with metrics into results log. | DUT (CM4) with eMMC writable; stress suite running; logging enabled; no requirement for external equipment. | PASS if (a) achieved sample rate error is ≤ 1% relative to expected 400 Hz and (b) simulator max lag ≤ 0.05 s. FAIL otherwise or on exceptions. |  |  |  |  | Ansh and Kranti
26 | HW–FW Integration | CM4 eMMC + OS file system + firmware test harness | eMMC Integrity Post-Stress | 1) After endurance and I/O stress complete, run the same checksum-based integrity procedure again on a new test file. 2) Write random data while computing SHA256, sync writes, read back, recompute SHA256, compare. 3) Record PASS/FAIL into results log. | DUT (CM4) with eMMC mounted; stress suite logging enabled. | PASS if post-stress read-back SHA256 matches the write-time SHA256 checksum. FAIL if mismatch or if read/write errors occur. |  |  |  |  | Ansh and Kranti
27 | HW–FW Integration | CM4 SDRAM + OS + memtester utility (if installed) | SDRAM Memtester | 1) Launch the SDRAM memtester test. 2) Run memtester using the configured memory size and loop count. 3) Capture return code and record PASS/FAIL into results log. | DUT (CM4); memtester installed and executable; logging enabled. | PASS if memtester returns exit code 0 and reports no errors. FAIL if memtester missing or returns non-zero exit code. |  |  |  |  | Ansh and Kranti
28 | HW–FW Integration | CM4 SDRAM + OS memory manager + stress-ng (if available) or Python allocator | SDRAM Memory Pressure | 1) Start memory pressure test for the defined duration. 2) If stress-ng is available, run memory pressure targeting ~90% memory for the test duration; otherwise allocate memory in Python in 1 MB chunks until duration completes (or MemoryError occurs). 3) Release allocated memory at end. 4) Record PASS/FAIL into results log. | DUT (CM4); stress-ng optional; logging enabled. | PASS if the pressure run completes to duration (stress-ng exit code 0), or (fallback) completes and a MemoryError is treated as acceptable evidence of high pressure. FAIL if the stress tool exits non-zero or other exceptions occur. |  |  |  |  | Ansh and Kranti
29 | HW–FW Integration | CM4 CPU + thermal sensor + OS telemetry + stress-ng (if available) | CPU Thermal Load (HW-CPU-001-A) | 1) Start CPU thermal test for the defined duration. 2) Apply sustained CPU load (stress-ng if present; otherwise CPU-burn threads). 3) At the defined log interval, record CPU temperature and CPU usage telemetry to the log. 4) Run until duration completes. 5) Record PASS/FAIL into results log. | DUT (CM4); logging enabled; stable environment for duration. | PASS if the test runs for full duration and logs periodic temperature/CPU usage without exception. (Throttling/overheat determination is made by reviewing the logs produced by the script.) |  |  |  |  | Ansh and Kranti
30 | HW–FW Integration | ECG/ICG acquisition simulator + CM4 CPU load + eMMC logging I/O + OS scheduler | Full Concurrency (HW-SDRAM-001-D / HW-CPU-001-B) | 1) Start full concurrency test for the defined duration. 2) Run the ECG/ICG simulator at 400 Hz writing binary stream files. 3) Run CPU burn concurrently. 4) Generate continuous disk logging bursts with forced flush/sync while also recording temperature/CPU telemetry. 5) Stop simulator and CPU load after duration. 6) Evaluate simulator rate and max lag vs schedule. 7) Record PASS/FAIL with metrics into results log. | DUT (CM4) with eMMC writable; logging enabled. | PASS if simulator rate error ≤ 1% and max lag ≤ 0.05 s under combined CPU + disk I/O load. FAIL otherwise or on exceptions. |  |  |  |  | Ansh and Kranti
31 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Apply MAX30009 measurement settings and observe state machine | 1) Connect to MAX30009 server. 2) Send a JSON request of type settings with measurement enabled and required parameters per the document. 3) Capture all asynchronous PUSH messages until the state machine indicates it is ready for polling. 4) Capture the actual_settings message. | DUT with MAX30009 service reachable (127.0.0.1:30009). Logging/packet capture enabled. | Firmware emits the documented meas_state sequence (pre_measuring → pre_measure_end → calibrating → calibrate_end → start_measuring) and then emits actual_settings. |  |  |  |  | Ansh and Kranti
32 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Invalid MAX30009 settings parameters handling (TBD) | 1) Connect to MAX30009 server. 2) Send a settings request with at least one parameter outside the documented allowed set (e.g., invalid string value for stimulate_current / invalid filter token). 3) Capture all responses and logs. | DUT with MAX30009 service reachable; capture logs and server responses. | Behavior not defined in the provided JSON documents; record firmware response(s)/log evidence and mark as spec clarification required. |  |  |  |  | Ansh and Kranti
33 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Poll MAX30009 get_data after start_measuring | 1) Ensure MAX30009 is in measurement mode (i.e., after start_measuring). 2) Send get_data request. 3) Capture the synchronous response payload. 4) Confirm the response includes a data array and includes sync-packet formatting where present. | DUT with MAX30009 configured and in measurement mode; capture responses. | Response type is data with fields including data_frequency, data_size, timestamp, data[]; sync-packet format is as documented for ICG [-999990000, sync_num, 0,0,0]. |  |  |  |  | Ansh and Kranti
34 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | get_data when not measuring returns no_measure | 1) Ensure MAX30009 is not in measurement mode (e.g., prior to start_measuring or after measurement is stopped/powered off). 2) Send get_data request. 3) Capture response. | DUT with MAX30009 service reachable; packet capture enabled. | Response is {"type":"no_measure"} when not in measurement mode. |  |  |  |  | Ansh and Kranti
35 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Start base-table build and observe asynchronous calib_data | 1) Connect to MAX30009 server. 2) Send build_base_table request. 3) Confirm immediate response. 4) Capture subsequent asynchronous calib_data JSON messages and store them. | DUT with MAX30009 service reachable; log capture enabled. | Immediate response is {"type":"build_base_table_started"}; asynchronous calib_data JSONs are emitted (per iteration) as documented. |  |  |  |  | Ansh and Kranti
36 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | build_base_table during active measurement (TBD) | 1) Put MAX30009 into measurement mode (per settings/meas_state). 2) While measuring, send build_base_table. 3) Capture response(s) and any emitted state/log messages. | DUT with MAX30009 in active measurement; capture logs + responses. | Behavior not defined in the provided documents for this ordering; record observed response(s)/state/log outputs for requirement clarification. |  |  |  |  | Ansh and Kranti
37 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Power off MAX30009 AFE | 1) Connect to MAX30009 server. 2) Send poweroff request. 3) Capture response. | DUT with MAX30009 service reachable. | Response is {"type":"power_is_off"}. |  |  |  |  | Ansh and Kranti
38 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | get_data after poweroff returns no_measure | 1) Send poweroff and confirm power_is_off. 2) Send get_data. 3) Capture response. | DUT with MAX30009 service reachable; capture responses. | After poweroff, MAX30009 is not in measurement mode; get_data returns {"type":"no_measure"} per documented behavior for non-measurement mode. |  |  |  |  | Ansh and Kranti
39 | FW-APP Integration | MAX30009 AFE + Firmware MAX30009 service | Repeat poweroff then re-enable measurement | 1) Send poweroff and confirm response. 2) Send settings with measurement enabled. 3) Capture meas_state sequence and actual_settings. | DUT with MAX30009 service reachable; capture logs/messages. | poweroff returns power_is_off; subsequent settings produces the documented meas_state sequence and actual_settings. |  |  |  |  | Ansh and Kranti
40 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | Apply ADS1293 sampling settings and enable conversion | 1) Connect to ADS1293 server. 2) Send settings with R-rates and conversion enable (per example). 3) Capture response. | DUT with ADS1293 service reachable (127.0.0.1:1293). | Firmware returns actual_settings reflecting the applied settings (as documented by example). |  |  |  |  | Ansh and Kranti
41 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | Invalid ADS1293 rate parameters handling (TBD) | 1) Connect to ADS1293 server. 2) Send settings with R2_rate or R3_rate outside documented allowed sets. 3) Capture response(s)/logs. | DUT with ADS1293 service reachable; log capture enabled. | Behavior not defined for invalid enum values in the provided reference; record observed response(s) for spec clarification. |  |  |  |  | Ansh and Kranti
42 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | Poll ADS1293 get_data and verify sync marker presence | 1) Ensure conversion is enabled via settings. 2) Send get_data. 3) Capture response payload. 4) Confirm sync marker format appears in the data stream. | DUT with ADS1293 configured and converting; capture responses. | Response type is data and includes sync marker entries with magic number -99999 and a sync counter as documented. |  |  |  |  | Ansh and Kranti
43 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | get_data before configuration/conversion (TBD) | 1) Start with ADS1293 service reachable but without sending settings to enable conversion. 2) Send get_data. 3) Capture response(s). | DUT with ADS1293 service reachable; capture responses. | Behavior not defined (no “no_measure” behavior specified for ADS1293); record observed response(s) for requirement clarification. |  |  |  |  | Ansh and Kranti
44 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | Power off ADS1293 AFE | 1) Connect to ADS1293 server. 2) Send poweroff. 3) Capture response. | DUT with ADS1293 service reachable. | Response is {"type":"power_is_off"}. |  |  |  |  | Ansh and Kranti
45 | FW-APP Integration | ADS1293 AFE + Firmware ADS1293 service | Repeat poweroff then re-enable conversion | 1) Send poweroff and confirm response. 2) Send settings to re-enable conversion. 3) Capture actual_settings. | DUT with ADS1293 service reachable; capture responses. | poweroff returns power_is_off; subsequent settings returns actual_settings. |  |  |  |  | Ansh and Kranti
48 | FW-APP Integration | Battery/charger telemetry + Firmware power service | Read battery information object | 1) Connect to power control server. 2) Send get_batt_info request. 3) Capture response payload. | DUT with power control service reachable (127.0.0.1:501). | Response is JSON object of type batt_info with expected fields (voltage, SOC, temperature, capacities, etc.) as shown. |  |  |  |  | Ansh and Kranti
49 | FW-APP Integration | Battery charging control + Firmware power service | Disable charging via JSON | 1) Connect to power control server. 2) Send charge_disable request. 3) Capture response. 4) Read get_batt_info and capture status fields indicating charge disable state. | DUT with power control service reachable; battery present. | Response is {"type":"charge_is_disable"}; subsequent batt_info reflects charge-disable state via reported status fields (e.g., battery_charge_is_disable). |  |  |  |  | Ansh and Kranti
50 | FW-APP Integration | Battery charging control + Firmware power service | Enable charging via JSON | 1) Connect to power control server. 2) Send charge_enable request. 3) Capture response. 4) Read get_batt_info and capture status fields indicating charge enable state. | DUT with power control service reachable; battery present. | Response is {"type":"charge_is_enable"}; subsequent batt_info reflects charge-enable state via reported status fields. |  |  |  |  | Ansh and Kranti
51 | FW-APP Integration | Firmware power service | Malformed/unknown power-control message handling (TBD) | 1) Connect to power control server. 2) Send an unsupported/unknown type. 3) Capture response/logs. | DUT with power control service reachable; capture logs. | Behavior not defined in provided power-control message examples; record observed behavior for spec clarification. |  |  |  |  | Ansh and Kranti
52 | FW-APP Integration | Button input + Firmware power/control reporting | Button_info telemetry increments hold_time while pressed | 1) Observe firmware output stream for button_info messages during a press-and-hold. 2) Capture messages from press start through release. | DUT with button accessible; logging enabled. | button_info messages are observed with state=true during hold and hold_time incrementing; upon release, state=false is reported (as shown). |  |  |  |  | Ansh and Kranti
53 | FW-APP Integration | ADS1293 + MAX30009 + Firmware sync insertion | Verify sync marker formats for ECG and ICG streams | 1) Configure both sensors for acquisition. 2) Poll both with get_data and capture packets containing sync marks. 3) Verify ECG sync format and ICG sync format (including scaling rules). | DUT with both servers reachable and data capture enabled. | ECG sync uses magic -99999 and sync_num as shown; ICG sync uses magic -999990000 and sync_num scaled by ×10000 (divide by 10000 to recover). |  |  |  |  | Ansh and Kranti
54 | FW-APP Integration | ADS1293 + MAX30009 + Firmware sync insertion | Sync counters monotonically increase every 1 second | 1) Start both sensors and begin polling via get_data. 2) Extract sync counters from each stream. 3) Verify ordering and 1-second progression per the documented validation description. | DUT with both services reachable; log extraction tooling. | Sync counters are monotonically increasing and represent 1-second sequence increments with no missing/duplicated sync frames (per the documented integration checks). |  |  |  |  | Ansh and Kranti
55 | FW-APP Integration | ADS1293 + MAX30009 + Firmware sync insertion | Common sync numbers align between streams within threshold | 1) Run both streams and collect timestamps for matching sync numbers from ECG and ICG. 2) Compute time deltas for each common sync mark. 3) Compare max delta to the documented threshold. | DUT with time-stamped capture of both streams. | Common sync marks exist; max time difference between ECG and ICG common sync marks is < 50 ms threshold per the provided pass criteria example. |  |  |  |  | Ansh and Kranti
56 | FW-APP Integration | MAX30009 + FW + resistive load | MAX30009 long-duration get_data at 0.5s polling | 1) Prepare per documented steps (including base table generation as described). 2) Configure MAX30009 for fs = 400 Hz and required settings per the described test. 3) Send get_data every 0.5 s for 1 hour and record. 4) Extract sync markers and compute mean sampling frequency from collected samples. 5) Power off at end. | DUT with MAX30009 connected to the described load; long-run capture enabled. | Sync markers are monotonically increasing every 1 s; mean sampling frequency over 1 hour is 400 Hz with 1 Hz error; poweroff returns power_is_off. |  |  |  |  | Ansh and Kranti
57 | FW-APP Integration | ADS1293 + FW + ECG simulator | ADS1293 60-second get_data at 0.5s polling | 1) Connect ECG simulator and set BPM values as described. 2) Configure ADS1293 for 400 Hz (R2=X, R3=Y per documented placeholder). 3) Send get_data every 0.5 s for 60 s. 4) Extract sync counters and compute effective sampling frequency from sample counts. 5) Power off. | DUT + ECG simulator; capture enabled. | SYNC_COUNTER values monotonically increase by 1 second with no missing frames during 60 s; mean sampling frequency is 400 Hz ±1 Hz; ECG signal integrity matches BPM; poweroff returns power_is_off. |  |  |  |  | Ansh and Kranti
58 | FW-APP Integration | ADS1293 + FW + ECG simulator | ADS1293 long-duration get_data at 0.5s polling | 1) Configure ADS1293 for 400 Hz as described. 2) Send get_data every 0.5 s for 1 hour and record. 3) Extract sync counters and compute effective sampling frequency over 1 hour. 4) Log sample counts, timestamps, sync counters, computed sampling frequency and signal quality metrics. 5) Power off. | DUT + ECG simulator; long-run capture enabled. | SYNC_COUNTER values monotonically increase every 1 second with no missing sync frames during 1 hour; mean sampling frequency is 400 Hz ±1 Hz; signal integrity matches BPM; poweroff returns power_is_off. |  |  |  |  | Ansh and Kranti
59 | FW-APP Integration | ADS1293 + MAX30009 + FW | Simultaneous acquisition + sync integrity + drift reporting | 1) Ensure both sensor servers are reachable. 2) Configure ADS1293 for 400 Hz and enable conversion. 3) Configure MAX30009 for 400 Hz per described settings and confirm meas_state sequence. 4) Send get_data to both servers every 0.5 s for 1 hour and record both streams. 5) Extract sync counters from both streams and verify monotonicity and no missing/duplicate frames. 6) Compute effective sampling frequency for both and compute drift using sync alignment. 7) Power off both sensors and verify responses. | DUT with ECG simulator + ICG resistive load as described; synchronized capture enabled. | Both streams show sync counters monotonically increasing every 1 second with no missing/duplicate sync frames over 1 hour; mean FS for both sensors is 400 Hz ±1 Hz; drift is computed and reported; poweroff returns power_is_off for both. |  |  |  |  | Ansh and Kranti
60 | FW-APP Integration | FW sensor services + client (mock app) | Reconnect behavior during repeated connect/disconnect | 1) Connect client to a sensor service port. 2) Open/close the connection repeatedly (5×) while the service is available. 3) Record whether reconnection is successful each cycle. | DUT; client capable of repeated connect/disconnect. | Reconnect succeeds each time and no socket leak is observed (as previously captured in the existing robustness criteria). |  |  |  |  | Ansh and Kranti
61 | FW-APP Integration | MAX30009 + FW + client | get_data polling before measurement is started | 1) Do not enable measurement. 2) Send get_data every 0.5 s for a defined short interval (e.g., 60 s) while remaining out of measurement mode. 3) Capture all responses. | DUT with MAX30009 server reachable. | Each response is no_measure while not in measurement mode. |  |  |  |  | Ansh and Kranti
62 | FW-APP Integration | ADS1293 + FW + client | ADS1293 get_data after poweroff | 1) Send poweroff to ADS1293 and confirm power_is_off. 2) Send get_data and capture response. | DUT with ADS1293 server reachable; capture responses. | Behavior not defined in provided ADS1293 examples for get_data-after-poweroff; record observed behavior for requirement clarification. |  |  |  |  | Ansh and Kranti
63 | FW-APP Integration | Storage (eMMC/SD) + base_table.json + Firmware MAX30009 service | Corrupted base_table.json on disk → service startup behavior and recovery evidence | 1) On the DUT storage, locate the MAX30009 base table file base_table.json. 2) Create a deliberate corruption in the file content (e.g., edit bytes/structure) while keeping a known-good backup available (as per the device’s backup mechanism). 3) Reboot the device (or restart the sensing service if the file is loaded at service init). 4) Start the MAX30009 sensing service and observe whether it reaches a ready state (service up and accepting connections). 5) Attempt to start measurement via the documented API flow (e.g., settings enabling measurement and observe the emitted service state messages). 6) Capture firmware logs and any observable service state messages during startup and during the first measurement attempt. | DUT with access to storage to corrupt base_table.json; ability to reboot device; ability to start the MAX30009 service and observe logs; client capable of connecting to the MAX30009 service and sending documented requests. | Firmware detects that the corrupted calibration/configuration file is bad via checksum and restores it from backup; objective evidence includes a log/event indicating restoration (e.g., “Corrupted file X restored from backup”) and the device/service proceeds using a valid file rather than failing silently. Record the observed MAX30009 service state during startup and measurement start attempt as the verification evidence. |  |  |  |  | Ansh and Kranti