## Test_Cases_Master_List_Final
Test_Case_ID | Test_Case_Name | Description | Component | Platform | Framework | Hardware_Required | Priority | Expected_Output | Pass_Criteria | Estimated_Duration | Service_Component
TC-PROTO-001 | TCP Connection ADS1293 Port | Verify TCP server accepts connection on ADS1293 service port 1293 | Protocol | CM4 Linux | pytest | No | CRITICAL | Connection established, socket returns success | Socket connection successful, no timeout | 30 sec | SPI_Service
TC-PROTO-002 | TCP Connection MAX30009 Port | Verify TCP server accepts connection on MAX30009 service port 30009 | Protocol | CM4 Linux | pytest | No | CRITICAL | Connection established, socket returns success | Socket connection successful, no timeout | 30 sec | SPI_Service
TC-PROTO-003 | TCP Connection WS2812 Port | Verify TCP server accepts connection on WS2812 LED service port 2812 | Protocol | CM4 Linux | pytest | No | HIGH | Connection established, socket returns success | Socket connection successful, no timeout | 30 sec | SPI_Service
TC-PROTO-004 | TCP Connection Power Service Port | Verify TCP server accepts connection on power control port 501 | Protocol | CM4 Linux | pytest | No | HIGH | Connection established, socket returns success | Socket connection successful, no timeout | 30 sec | Power_Service
TC-PROTO-005 | Valid JSON Start Command ECG | Send valid start command for ECG streaming and verify acknowledgment | Protocol | CM4 Linux | pytest | No | CRITICAL | JSON response: {"status":"ok","cmd":"start","sensor":"ecg"} | Status = ok, command echoed correctly | 1 min | JSON_Protocol + SPI_Service
TC-PROTO-006 | Valid JSON Stop Command | Send valid stop command and verify streaming stops | Protocol | CM4 Linux | pytest | No | CRITICAL | JSON response: {"status":"ok","cmd":"stop"} | Streaming stopped, status = ok | 1 min | JSON_Protocol
TC-PROTO-007 | Valid JSON Configure Command | Configure sensor parameters via JSON | Protocol | CM4 Linux | pytest | No | HIGH | JSON response: {"status":"ok","cmd":"configure"} | Configuration applied, ack received | 1 min | JSON_Protocol
TC-PROTO-008 | Valid JSON Status Query | Query current system status via JSON | Protocol | CM4 Linux | pytest | No | HIGH | JSON with status fields: streaming state, config, errors | Valid JSON returned with all status fields | 30 sec | JSON_Protocol
TC-PROTO-009 | Malformed JSON Handling | Server handles malformed JSON gracefully without crash | Protocol | CM4 Linux | pytest | No | CRITICAL | JSON response: {"status":"error","reason":"invalid_json"} | Error returned, server remains stable | 1 min | JSON_Protocol
TC-PROTO-010 | Missing Required Field in JSON | Server rejects commands with missing required fields | Protocol | CM4 Linux | pytest | No | HIGH | JSON response: {"status":"error","reason":"missing_field"} | Error message identifies missing field | 30 sec | JSON_Protocol
TC-PROTO-011 | Unknown Command Handling | Server handles unknown commands without crash | Protocol | CM4 Linux | pytest | No | MEDIUM | JSON response: {"status":"error","reason":"unknown_command"} | Error returned, server continues running | 30 sec | JSON_Protocol
TC-PROTO-012 | Partial Frame Reception | Server handles fragmented JSON frames correctly | Protocol | CM4 Linux | pytest | No | MEDIUM | Server waits for complete frame or timeout, then processes | Complete frame processed or timeout error returned | 1 min | JSON_Protocol
TC-PROTO-013 | Oversized JSON Payload | Server rejects JSON payloads exceeding size limit | Protocol | CM4 Linux | pytest | No | MEDIUM | JSON response: {"status":"error","reason":"payload_too_large"} | Oversized payload rejected, no buffer overflow | 1 min | JSON_Protocol
TC-PROTO-014 | Concurrent Client Connections | Multiple clients connect simultaneously without interference | Protocol | CM4 Linux | pytest | No | HIGH | All clients receive valid responses to their requests | 4 concurrent clients handled correctly | 2 min | JSON_Protocol
TC-PROTO-015 | Rapid Connect Disconnect | Server handles connection churn without resource leaks | Protocol | CM4 Linux | pytest | No | MEDIUM | 100 connect/disconnect cycles complete successfully | No resource leaks, all connections handled | 3 min | JSON_Protocol
TC-PROTO-016 | Sequential Command Flow | Execute configure->’start->status->stop sequence | Protocol | CM4 Linux | pytest | No | CRITICAL | All commands return status ok in sequence | Complete workflow executes successfully | 2 min | JSON_Protocol
TC-PROTO-017 | Command Response Timing | Measure TCP response latency for status command | Protocol | CM4 Linux | pytest | No | MEDIUM | Response time measured and logged | Response time < 100ms | 1 min | JSON_Protocol
TC-PROTO-018 | LED Color Command | Send LED color control command via JSON | Protocol | CM4 Linux | pytest | No | HIGH | JSON response: {"status":"ok","cmd":"led"} | Command accepted, ack received | 30 sec | JSON_Protocol + SPI_Service
TC-PROTO-019 | Power Control Command | Send power control command to toggle power rail | Protocol | CM4 Linux | pytest | No | HIGH | JSON response: {"status":"ok","cmd":"power"} | Power command accepted | 30 sec | JSON_Protocol
TC-PROTO-020 | Battery Query Command | Query battery status via power service | Protocol | CM4 Linux | pytest | No | HIGH | JSON with battery data: voltage, current, SoC, capacity | Valid battery telemetry returned | 1 min | JSON_Protocol + Power_Service
TC-PROTO-021 | Impedance Measurement Command | Request bioimpedance measurement via MAX30009 service | Protocol | CM4 Linux | pytest | No | HIGH | JSON response: {"status":"ok","cmd":"measure"} | Measurement command accepted | 30 sec | JSON_Protocol + SPI_Service
TC-PROTO-022 | Read Configuration Command | Read current sensor configuration via JSON | Protocol | CM4 Linux | pytest | No | MEDIUM | JSON with current config: rate, gain, channels | Config values match what was set | 1 min | JSON_Protocol
TC-PROTO-023 | Conflicting Concurrent Commands | Handle conflicting commands from multiple clients | Protocol | CM4 Linux | pytest | No | MEDIUM | Second command queued or rejected with error | Conflict resolved without crash | 2 min | JSON_Protocol
TC-PROTO-024 | Idempotent Start Command | Starting already-started stream returns appropriate error | Protocol | CM4 Linux | pytest | No | MEDIUM | JSON response: {"status":"error","reason":"already_streaming"} | Duplicate start detected and rejected | 30 sec | JSON_Protocol
TC-PROTO-025 | Stop Before Start | Stopping non-streaming sensor handled gracefully | Protocol | CM4 Linux | pytest | No | LOW | JSON response: ok or error with appropriate reason | Handled without crash | 30 sec | JSON_Protocol + SPI_Service
TC-DRV-GPIO-001 | GPIO Driver Initialization | Initialize GPIO driver successfully for valid pin number | Drivers | Laptop/CM4 | gtest | No | CRITICAL | MockGPIODriver created, is_valid() returns true | GPIO driver initialized for pin 17 | 10 sec | Infrastructure
TC-DRV-GPIO-002 | GPIO Set Output Direction | Configure GPIO pin as output | Drivers | Laptop/CM4 | gtest | No | HIGH | Direction set to VT_GPIO_OUTPUT | get_direction() returns VT_GPIO_OUTPUT | 10 sec | Infrastructure
TC-DRV-GPIO-003 | GPIO Write High State | Write HIGH to GPIO output pin | Drivers | Laptop/CM4 | gtest | No | HIGH | GPIO state set to HIGH | get_GPIO() returns true | 10 sec | Infrastructure
TC-DRV-GPIO-004 | GPIO Write Low State | Write LOW to GPIO output pin | Drivers | Laptop/CM4 | gtest | No | HIGH | GPIO state set to LOW | get_GPIO() returns false | 10 sec | Infrastructure
TC-DRV-GPIO-005 | GPIO Toggle Operations | Toggle GPIO pin multiple times between HIGH and LOW | Drivers | Laptop/CM4 | gtest | No | MEDIUM | Pin toggles: LOW->HIGH->LOW->HIGH | All state changes successful | 20 sec | Infrastructure
TC-DRV-GPIO-006 | GPIO Set Input Direction | Configure GPIO pin as input | Drivers | Laptop/CM4 | gtest | No | HIGH | Direction set to VT_GPIO_INPUT | get_direction() returns VT_GPIO_INPUT | 10 sec | Infrastructure
TC-DRV-GPIO-007 | GPIO Cannot Write to Input | Writing to input-configured pin fails appropriately | Drivers | Laptop/CM4 | gtest | No | MEDIUM | set_GPIO() returns false | Write operation rejected for input pin | 10 sec | Infrastructure
TC-DRV-GPIO-008 | GPIO Release Pin | Release GPIO pin resources | Drivers | Laptop/CM4 | gtest | No | MEDIUM | Direction set to VT_GPIO_RELEASE, state cleared | Pin released successfully | 10 sec | Infrastructure
TC-DRV-GPIO-009 | GPIO Invalid Pin Number | Reject invalid GPIO pin numbers | Drivers | Laptop/CM4 | gtest | No | MEDIUM | is_valid() returns false for pin > 27 | Invalid pin number detected | 10 sec | Infrastructure
TC-DRV-GPIO-010 | Multiple GPIO Independence | Multiple GPIO instances operate independently | Drivers | Laptop/CM4 | gtest | No | MEDIUM | Pin 17 HIGH, Pin 18 LOW maintained independently | No crosstalk between GPIO instances | 30 sec | Infrastructure
TC-DRV-SPI-001 | SPI Device Open | Open SPI device successfully | Drivers | Laptop/CM4 | gtest | No | CRITICAL | SPI device /dev/spidev0.0 opened | open() returns true | 10 sec | SPI_Service
TC-DRV-SPI-002 | SPI Transfer Operation | Perform SPI read/write transaction | Drivers | Laptop/CM4 | gtest | No | HIGH | RX buffer filled with data | RX buffer size == TX buffer size | 20 sec | SPI_Service
TC-DRV-SPI-003 | SPI Configuration | Configure SPI mode and speed correctly | Drivers | Laptop/CM4 | gtest | No | MEDIUM | Mode set to MODE_0, speed set to 5MHz | Configuration accepted | 15 sec | SPI_Service
TC-DRV-I2C-001 | I2C Device Open | Open I2C device successfully | Drivers | Laptop/CM4 | gtest | No | HIGH | I2C device /dev/i2c-1 opened with address 0x0B | open() returns true | 10 sec | SPI_Service
TC-DRV-I2C-002 | I2C Read Operation | Read byte from I2C device register | Drivers | Laptop/CM4 | gtest | No | HIGH | Valid data byte returned from register read | Data returned matches expected | 15 sec | SPI_Service
TC-DRV-I2C-003 | I2C Write Operation | Write byte to I2C device register | Drivers | Laptop/CM4 | gtest | No | HIGH | Write operation completes successfully | write_byte() returns true | 15 sec | SPI_Service
TC-DRV-I2C-004 | ADS1293 Register Access | Access ADS1293 registers via SPI mock | Drivers | Laptop/CM4 | gtest | No | HIGH | Register read/write operations execute correctly | Valid register access pattern | 20 sec | SPI_Service
TC-DRV-I2C-005 | MAX30009 Register Access | Access MAX30009 registers via I2C mock | Drivers | Laptop/CM4 | gtest | No | HIGH | Register read/write operations execute correctly | Valid register access pattern | 20 sec | SPI_Service
TC-DATA-001 | ECG FIFO Buffer Write | Write ECG samples to circular FIFO buffer | Data Processing | Laptop/CM4 | gtest | No | HIGH | 10 samples written to FIFO | _IFIFO_write_pos increments to 10 | 20 sec | SPI_Service
TC-DATA-002 | ECG FIFO Buffer Read | Read ECG samples from FIFO buffer | Data Processing | Laptop/CM4 | gtest | No | HIGH | 10 samples read from FIFO | Read data matches written data | 20 sec | SPI_Service
TC-DATA-003 | FIFO Buffer Wraparound | FIFO circular buffer wraps around correctly | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | Write position wraps from 3000 to 100 after 3100 writes | Circular wraparound successful | 30 sec | SPI_Service
TC-DATA-004 | FIFO Buffer Full Detection | Detect when FIFO reaches capacity | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | Full condition detected at 3000 samples | Full flag set or overflow handled | 20 sec | SPI_Service
TC-DATA-005 | FIFO Buffer Empty Detection | Detect when FIFO is empty | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | Empty condition detected when read_pos == write_pos | Empty condition returned | 10 sec | SPI_Service
TC-DATA-006 | Sync Mark Insertion | Insert sync mark (-99999) into data stream | Data Processing | Laptop/CM4 | gtest | No | HIGH | SYNC_MARK_MAGIC_NUM (-99999) inserted at correct position | Sync mark appears in FIFO at expected location | 15 sec | SPI_Service
TC-DATA-007 | Sample Rate Calculation | Calculate actual sampling rate from collected data | Data Processing | Laptop/CM4 | gtest | No | HIGH | Calculated rate within Â±5% of configured rate | Rate = samples/time â‰ˆ configured rate | 30 sec | SPI_Service
TC-DATA-008 | Data Format Conversion | Convert raw ADC values to voltage | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | Voltage value calculated correctly from raw ADC | Voltage within expected range | 15 sec | SPI_Service
TC-DATA-009 | Multi-Channel Data Interleaving | Handle 3-channel ECG data without mixing channels | Data Processing | Laptop/CM4 | gtest | No | HIGH | Channel 1, 2, 3 data stored separately and correctly | Channels maintain distinct data | 30 sec | SPI_Service
TC-DATA-010 | Impedance Magnitude Calculation | Calculate impedance magnitude from I/Q data | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | For I=100, Q=100: magnitude = sqrt(100Â²+100Â²) = 141.42 | Calculated magnitude â‰ˆ 141.42 | 10 sec | SPI_Service
TC-DATA-011 | Impedance Phase Calculation | Calculate impedance phase from I/Q data | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | For I=100, Q=100: phase = atan2(100,100) = 0.785 rad (45Â°) | Phase angle â‰ˆ 45Â° | 10 sec | SPI_Service
TC-DATA-012 | Timestamp Generation | Generate ISO8601 timestamps for samples | Data Processing | Laptop/CM4 | gtest | No | LOW | Timestamp string in format YYYY-MM-DDTHH:MM:SS | Valid ISO8601 timestamp | 10 sec | SPI_Service
TC-DATA-013 | Data Decimation | Decimate data by factor of 2 | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | 100 input samples â†’ 50 output samples | Output count = input count / decimation factor | 15 sec | SPI_Service
TC-DATA-014 | JSON Data Serialization | Convert sample data to JSON format | Data Processing | Laptop/CM4 | gtest | No | HIGH | Valid JSON with all sample fields | JSON parseable with all required fields | 20 sec | SPI_Service
TC-DATA-015 | Circular Buffer Thread Safety | Verify FIFO is thread-safe during concurrent access | Data Processing | Laptop/CM4 | gtest | No | MEDIUM | No data corruption during concurrent read/write | All samples read correctly, no corruption | 1 min | SPI_Service
TC-STATE-001 | State Transition Idle to Streaming | Transition from idle state to streaming state | State Machine | CM4 Linux | pytest | No | CRITICAL | State changes from idle to streaming | Status query shows streaming=true | 30 sec | JSON_Protocol + SPI_Service
TC-STATE-002 | State Transition Streaming to Stopped | Transition from streaming to stopped state | State Machine | CM4 Linux | pytest | No | CRITICAL | Streaming stops, state returns to idle | Status query shows streaming=false | 30 sec | JSON_Protocol + SPI_Service
TC-STATE-003 | State Transition Configure Then Start | Configure sensor then start streaming | State Machine | CM4 Linux | pytest | No | HIGH | Configured parameters applied, streaming starts | Configuration active during streaming | 1 min | JSON_Protocol + SPI_Service
TC-STATE-004 | Invalid Transition Configure During Streaming | Cannot reconfigure while streaming | State Machine | CM4 Linux | pytest | No | MEDIUM | Error response: cannot configure while streaming | Configure command rejected during streaming | 30 sec | JSON_Protocol + SPI_Service
TC-STATE-005 | Power State Enable ADS1293 | Enable power rail for ADS1293 sensor | State Machine | CM4 Linux | pytest | No | HIGH | Power rail enabled, GPIO HIGH | Power enable acknowledged | 30 sec | JSON_Protocol + Power_Service
TC-STATE-006 | Power State Disable ADS1293 | Disable power rail for ADS1293 sensor | State Machine | CM4 Linux | pytest | No | HIGH | Power rail disabled, GPIO LOW, streaming stops | Power disable successful, streaming terminated | 30 sec | JSON_Protocol + Power_Service
TC-STATE-007 | LED State Solid Color | Set LED strip to solid color | State Machine | CM4 Linux | pytest | No | MEDIUM | LED command accepted | Command acknowledged | 30 sec | JSON_Protocol + SPI_Service
TC-STATE-008 | LED State Blink Pattern | Set LED strip to blink pattern | State Machine | CM4 Linux | pytest | No | MEDIUM | Blink pattern command accepted | Pattern command acknowledged | 30 sec | JSON_Protocol + SPI_Service
TC-STATE-009 | Error State Sensor Not Responding | Handle sensor communication failure | State Machine | CM4 Linux | pytest | No | HIGH | Error state entered, error message returned | Sensor timeout detected and reported | 1 min | JSON_Protocol + SPI_Service
TC-STATE-010 | Error State Recovery Automatic Retry | System retries failed sensor operations | State Machine | CM4 Linux | pytest | No | MEDIUM | Retry attempts made, eventual success or final error | Automatic retry mechanism functions | 2 min | JSON_Protocol + SPI_Service
TC-STATE-011 | Workflow Complete Acquisition Cycle | Execute full acquisition workflow from configure to data collection | State Machine | CM4 Linux | pytest | No | CRITICAL | 5000 samples collected (500Hz Ã— 10s), data file created | Complete workflow successful, expected sample count | 2 min | JSON_Protocol + SPI_Service
TC-STATE-012 | Workflow Multi-Sensor Acquisition | Acquire from ADS1293 and MAX30009 simultaneously | State Machine | CM4 Linux | pytest | No | HIGH | Both sensors streaming concurrently, data from both | Dual sensor acquisition successful | 2 min | JSON_Protocol + SPI_Service
TC-ERR-001 | Null Pointer Handling | Handle null pointer passed to driver function | Error Handling | Laptop/CM4 | gtest | No | HIGH | Function returns error, no crash or segfault | Error returned, no crash | 10 sec | SPI_Service
TC-ERR-002 | Out of Bounds Array Access | Prevent buffer overflow in FIFO access | Error Handling | Laptop/CM4 | gtest | No | HIGH | Bounds check prevents access, error returned | Out of bounds access blocked | 10 sec | SPI_Service
TC-ERR-003 | Division by Zero Protection | Protect against division by zero in calculations | Error Handling | Laptop/CM4 | gtest | No | MEDIUM | Returns default value or error, no crash | Division by zero handled gracefully | 10 sec | SPI_Service
TC-ERR-004 | TCP Connection Timeout | Handle TCP connection timeout | Error Handling | CM4 Linux | pytest | No | MEDIUM | TimeoutError raised after 3 seconds | Timeout detected and raised | 10 sec | JSON_Protocol + SPI_Service
TC-ERR-005 | TCP Connection Lost During Streaming | Handle abrupt connection loss during streaming | Error Handling | CM4 Linux | pytest | No | HIGH | Server cleans up connection, continues running | Connection cleaned up, server stable | 1 min | SPI_Service
TC-ERR-006 | Invalid Parameter Rate Too High | Reject sampling rate exceeding maximum | Error Handling | CM4 Linux | pytest | No | MEDIUM | Error: invalid_rate | Out-of-range rate rejected | 30 sec | SPI_Service
TC-ERR-007 | Invalid Parameter Rate Too Low | Reject sampling rate below minimum | Error Handling | CM4 Linux | pytest | No | MEDIUM | Error: invalid_rate | Out-of-range rate rejected | 30 sec | SPI_Service
TC-ERR-008 | Negative Duration Value | Reject negative duration parameter | Error Handling | CM4 Linux | pytest | No | LOW | Error: invalid_duration | Negative duration rejected | 30 sec | SPI_Service
TC-ERR-009 | Missing Sensor Field in Command | Handle missing required sensor field in JSON | Error Handling | CM4 Linux | pytest | No | MEDIUM | Error: missing_sensor_field | Missing field detected and reported | 30 sec | JSON_Protocol + SPI_Service
TC-ERR-010 | Unknown Sensor Type | Handle unknown sensor name in command | Error Handling | CM4 Linux | pytest | No | LOW | Error: unknown_sensor | Unknown sensor rejected | 30 sec | SPI_Service
TC-CALIB-001 | Load Calibration File | Load calibration JSON file successfully | Calibration | CM4 Linux | pytest | No | HIGH | Calibration data structure populated from calib/0_0.json | Calibration file loaded successfully | 30 sec | SPI_Service
TC-CALIB-002 | Missing Calibration File | Handle missing calibration file gracefully | Calibration | CM4 Linux | pytest | No | MEDIUM | Error returned, default calibration used | Missing file handled, defaults applied | 30 sec | SPI_Service
TC-CALIB-003 | Corrupted Calibration File | Handle corrupted JSON in calibration file | Calibration | CM4 Linux | pytest | No | MEDIUM | Parse error detected, default calibration used | Corruption detected, defaults used | 30 sec | SPI_Service
TC-CALIB-004 | Validate Calibration Data Range | Verify calibration values within expected ranges | Calibration | CM4 Linux | pytest | No | MEDIUM | All 85 files validated: magnitude > 0, phase -180Â° to +180Â° | All calibration values valid | 5 min | SPI_Service
TC-CALIB-005 | Save Calibration Data | Save new calibration data to JSON file | Calibration | CM4 Linux | pytest | No | HIGH | File created with calibration data, reload matches | Calibration saved and verified | 1 min | SPI_Service
TC-CALIB-006 | Calibration Timestamp Format | Verify calibration timestamp is ISO8601 format | Calibration | CM4 Linux | pytest | No | LOW | Timestamp in format YYYY-MM-DDTHH:MM:SS | Valid timestamp format | 10 sec | SPI_Service
TC-CALIB-007 | Apply Calibration to Raw Data | Apply gain correction to raw impedance measurement | Calibration | Laptop/CM4 | gtest | No | MEDIUM | Raw=100Î©, gain=1.05 â†’ corrected=105Î© | Calibration applied correctly | 15 sec | SPI_Service
TC-CALIB-008 | Calibration File Naming Convention | Verify calibration files follow naming convention | Calibration | CM4 Linux | pytest | No | LOW | All files match pattern {current_index}_{frequency_index}.json | Naming convention followed | 1 min | SPI_Service
TC-SENSOR-ADS-001 | ADS1293 Device ID Read | Read and verify ADS1293 chip ID on hardware | ADS1293 | CM4 Linux | pytest | Yes - PCB | CRITICAL | Device ID = 0xAD (or per datasheet) | Correct device ID read | 1 min | SPI_Service
TC-SENSOR-ADS-002 | ADS1293 ECG Data Acquisition | Acquire real ECG samples from hardware | ADS1293 | CM4 Linux | pytest | Yes - PCB | CRITICAL | 2500 samples collected (250Hz Ã— 10s) Â±2% | Expected sample count acquired | 2 min | SPI_Service
TC-SENSOR-ADS-003 | ECG Signal R-Peak Detection | Detect R-peaks in acquired ECG waveform | ADS1293 | CM4 Linux | pytest | Yes - PCB | HIGH | R-peaks detected, HR calculated between 40-180 bpm | Valid heart rate calculated | 3 min | SPI_Service
TC-SENSOR-ADS-004 | ECG Signal Noise Floor | Verify ECG signal noise is acceptable | ADS1293 | CM4 Linux | pytest | Yes - PCB | MEDIUM | Standard deviation > 0.5 (non-flat signal) | Signal variance above threshold | 2 min | SPI_Service
TC-SENSOR-ADS-005 | ECG Sampling Rate Accuracy | Verify actual sampling rate matches configured rate | ADS1293 | CM4 Linux | pytest | Yes - PCB | Actual rate = 500Hz Â±5% (475-525 Hz) | Sampling rate accurate | 2 min |  | JSON_Protocol + SPI_Service
TC-SENSOR-ADS-006 | ADS1293 Multi-Channel Acquisition | Acquire from all 3 ECG channels simultaneously | ADS1293 | CM4 Linux | pytest | Yes - PCB | MEDIUM | 300 samples (3 channels Ã— 100), channels distinct | All 3 channels acquired correctly | 2 min | SPI_Service
TC-SENSOR-ADS-007 | ADS1293 Leads-Off Detection | Detect electrode disconnection | ADS1293 | CM4 Linux | pytest | Yes - PCB | HIGH | Leads-off status=true when electrode disconnected | Leads-off detected correctly | 2 min | SPI_Service
TC-SENSOR-ADS-008 | ADS1293 Long Duration Stability | Stream ECG continuously for 1 hour | ADS1293 | CM4 Linux | pytest | Yes - PCB | HIGH | < 0.1% packet loss, no FIFO overflow | 1 hour stable streaming | 60 min | SPI_Service
TC-SENSOR-ADS-009 | ADS1293 Power Cycle Recovery | Verify sensor recovers after power cycle | ADS1293 | CM4 Linux | pytest | Yes - PCB | HIGH | Sensor re-initializes and resumes streaming after power cycle | Power cycle recovery successful | 5 min | SPI_Service
TC-SENSOR-MAX-001 | MAX30009 Chip ID Read | Read and verify MAX30009 chip ID on hardware | MAX30009 | CM4 Linux | pytest | Yes - PCB | CRITICAL | Chip ID matches datasheet value | Correct chip ID read | 1 min | SPI_Service
TC-SENSOR-MAX-002 | MAX30009 Impedance Measurement | Measure bioimpedance on hardware | MAX30009 | CM4 Linux | pytest | Yes - PCB | CRITICAL | Impedance value 50-2000Î© (typical range) | Valid impedance measured | 2 min | SPI_Service
TC-SENSOR-MAX-003 | Impedance Stability Test | Verify impedance readings stable over time | MAX30009 | CM4 Linux | pytest | Yes - PCB | HIGH | Coefficient of variation < 20% over 100 readings | Impedance readings stable | 3 min | SPI_Service
TC-SENSOR-MAX-004 | Impedance Frequency Sweep | Measure impedance across frequency range | MAX30009 | CM4 Linux | pytest | Yes - PCB | MEDIUM | Valid impedance at 1kHz, 5kHz, 10kHz, 50kHz, 100kHz | Frequency sweep successful | 5 min | SPI_Service
TC-SENSOR-MAX-005 | MAX30009 Interrupt Timing | Verify DRDY interrupt timing accuracy | MAX30009 | CM4 Linux | pytest | Yes - PCB | MEDIUM | Interrupt period matches configured rate (e.g., 4ms for 250Hz) Â±0.2ms | Interrupt timing accurate | 2 min | SPI_Service
TC-SENSOR-MAX-006 | MAX30009 Calibration Application | Apply calibration and verify accuracy with known resistor | MAX30009 | CM4 Linux | pytest | Yes - PCB + Resistor | HIGH | Measured 1000Î© resistor reads 950-1050Î© (Â±5%) | Calibration applied accurately | 3 min | SPI_Service
TC-SENSOR-MAX-007 | MAX30009 Long Duration Stability | Measure impedance continuously for 1 hour | MAX30009 | CM4 Linux | pytest | Yes - PCB | MEDIUM | < 0.1% packet loss, mean stable (CV < 10%) | 1 hour stable measurement | 60 min | SPI_Service
TC-LED-001 | WS2812 Initialization | Initialize WS2812 LED strip | WS2812 | CM4 Linux | pytest | Yes - PCB | HIGH | Initialization successful, no errors | LED strip initialized | 30 sec | SPI_Service
TC-LED-002 | WS2812 Solid Red Color | Set all LEDs to solid red color | WS2812 | CM4 Linux | manual | Yes - PCB | HIGH | All LEDs display red (visual verification) | Red color displayed | 30 sec | SPI_Service
TC-LED-003 | WS2812 Color Sequence | Cycle through red â†’ green â†’ blue sequence | WS2812 | CM4 Linux | manual | Yes - PCB | MEDIUM | LEDs change colors in sequence (visual) | Color sequence correct | 1 min | SPI_Service
TC-LED-004 | WS2812 Timing Validation | Verify WS2812 protocol timing with logic analyzer | WS2812 | CM4 Linux | pytest | Yes - PCB + Analyzer | LOW | T0H=350ns, T1H=700ns, reset=50Âµs (all Â±150ns) | Timing within spec | 5 min | SPI_Service
TC-POWER-001 | Battery Voltage Reading | Read battery voltage via SMBus | Power Control | CM4 Linux | pytest | Yes - PCB | HIGH | Voltage reading 3.0-4.2V (Li-ion range) | Valid voltage read | 1 min | Power_Service
TC-POWER-002 | Battery Current Reading | Read battery current via SMBus | Power Control | CM4 Linux | pytest | Yes - PCB | MEDIUM | Current reading 0-2000mA (reasonable range) | Valid current read | 1 min | Power_Service
TC-POWER-003 | Battery State of Charge | Read battery SoC percentage | Power Control | CM4 Linux | pytest | Yes - PCB | MEDIUM | SoC 0-100% | Valid SoC percentage | 1 min | Power_Service
TC-POWER-004 | Power Rail Enable | Enable power rail via GPIO | Power Control | CM4 Linux | pytest | Yes - PCB | HIGH | GPIO HIGH, power rail enabled | Power rail enabled | 30 sec | Power_Service
TC-POWER-005 | Power Rail Disable | Disable power rail via GPIO | Power Control | CM4 Linux | pytest | Yes - PCB | HIGH | GPIO LOW, power rail disabled | Power rail disabled | 30 sec | Power_Service
TC-POWER-006 | Low Battery Warning | System detects low battery condition | Power Control | CM4 Linux | pytest | Yes - PCB | MEDIUM | Warning message generated when voltage < 3.3V | Low battery detected | 2 min | Power_Service
TC-STRESS-001 | 1 Hour Continuous ECG Streaming | Stream ECG data continuously for 1 hour | Stress Test | CM4 Linux | pytest | Yes - PCB | HIGH | < 0.1% packet loss, no crashes, 900000 samples (250Hz Ã— 3600s) | 1 hour stable streaming | 60 min | SPI_Service
TC-STRESS-002 | Memory Leak Detection | Detect memory leaks during long operation | Stress Test | CM4 Linux | pytest | Yes - PCB | HIGH | Memory usage stable Â±10MB over 2 hours | No memory leaks detected | 120 min | SPI_Service
TC-STRESS-003 | CPU Usage Under Load | Monitor CPU usage during multi-sensor streaming | Stress Test | CM4 Linux | pytest | Yes - PCB | MEDIUM | Average CPU < 50%, peak CPU < 80% | CPU usage acceptable | 30 min | SPI_Service
TC-STRESS-004 | Rapid Start Stop Cycles | Stress test start/stop operations | Stress Test | CM4 Linux | pytest | Yes - PCB | MEDIUM | 1000 start/stop cycles complete successfully | All cycles successful | 25 min | SPI_Service
TC-STRESS-005 | High Frequency JSON Commands | Send JSON commands at high rate | Stress Test | CM4 Linux | pytest | No | MEDIUM | 1000 rapid status queries all answered, no timeouts | All requests handled | 5 min | JSON_Protocol + SPI_Service
TC-STRESS-006 | Filesystem Stress Continuous Logging | Verify filesystem handles continuous data writes | Stress Test | CM4 Linux | pytest | Yes - PCB | MEDIUM | 2 hours continuous logging, no disk errors, file valid | Continuous logging successful | 120 min | SPI_Service
TC-MAIN-001 | Application Startup | Launch main application and verify all services start | Main Application | CM4 Linux | pytest | No | CRITICAL | All 3 TCP servers listening on ports 1293, 30009, 2812 | All services started | 1 min | JSON_Protocol + Power_Service + SPI_Service
TC-MAIN-002 | Application Shutdown | Clean shutdown of main application | Main Application | CM4 Linux | pytest | No | CRITICAL | All resources freed, graceful exit | Clean shutdown successful | 1 min | JSON_Protocol + Power_Service + SPI_Service
TC-MAIN-003 | Application Crash Recovery | Verify application recovers after crash | Main Application | CM4 Linux | pytest | No | HIGH | Application restarts, all services functional | Crash recovery successful | 2 min | JSON_Protocol + Power_Service + SPI_Service
TC-INT-001 | Simultaneous Multi-Sensor Acquisition | Run ECG + Impedance + LED simultaneously | System Integration | CM4 Linux | pytest | Yes - PCB | HIGH | All 3 sensors operating without interference | Concurrent operation successful | 5 min | JSON_Protocol + Power_Service + SPI_Service
TC-INT-002 | Cross-Service Communication | LED changes based on ECG heart rate | System Integration | CM4 Linux | pytest | Yes - PCB | MEDIUM | LED color changes when HR threshold crossed | Service coordination successful | 3 min | JSON_Protocol + Power_Service + SPI_Service
TC-INT-003 | System Resource Usage | Monitor system resources during full load | System Integration | CM4 Linux | pytest | Yes - PCB | MEDIUM | System responsive, CPU < 70%, memory stable | System remains stable under load | 10 min | JSON_Protocol + Power_Service + SPI_Service
TC-MAX-LIB-001 | MAX30009 I2C Initialization | Initialize I2C communication with MAX30009 | MAX30009 Library | Laptop/CM4 | gtest | No | CRITICAL | I2C device opened at address 0x0B | I2C initialized successfully | 10 sec | SPI_Service
TC-MAX-LIB-002 | MAX30009 Register Read | Read register from MAX30009 via I2C mock | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | Correct I2C transaction format, value returned | Register read successful | 15 sec | SPI_Service
TC-MAX-LIB-003 | MAX30009 Register Write | Write to MAX30009 register via I2C mock | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | Correct I2C write transaction, readback verifies | Register write successful | 15 sec | SPI_Service
TC-MAX-LIB-004 | MAX30009 Frequency Configuration | Configure measurement frequency (1kHz-100kHz) | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | Correct register values for 1kHz, 10kHz, 50kHz, 100kHz | Frequency configured correctly | 20 sec | JSON_Protocol + SPI_Service
TC-MAX-LIB-005 | MAX30009 Current Selection | Set excitation current (5 levels) | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | Correct DAC register values for current index 0-4 | Current level set correctly | 15 sec | SPI_Service
TC-MAX-LIB-006 | MAX30009 Gain Configuration | Set ADC gain (1x,2x,4x,8x) | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | All gain settings accepted | All 4 gain configurations successful |  | SPI_Service
TC-MAX-LIB-007 | MAX30009 Start Measurement | Start impedance measurement | MAX30009 Library | Laptop/CM4 | gtest | No | CRITICAL | ADC enabled, clock started, FIFO cleared | Measurement started | 15 sec | SPI_Service
TC-MAX-LIB-008 | MAX30009 Stop Measurement | Stop impedance measurement | MAX30009 Library | Laptop/CM4 | gtest | No | CRITICAL | ADC disabled, FIFO data preserved | Measurement stopped | 15 sec | SPI_Service
TC-MAX-LIB-009 | MAX30009 Data Ready Check | Poll data ready status | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | Status register read, DRDY bit checked | Data ready status correct | 10 sec | SPI_Service
TC-MAX-LIB-010 | MAX30009 I-Channel Read | Read in-phase (I) channel data | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | 24-bit signed value read, converted to int32_t | I-channel read successfully | 15 sec | SPI_Service
TC-MAX-LIB-011 | MAX30009 Q-Channel Read | Read quadrature (Q) channel data | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | 24-bit signed value read, converted to int32_t | Q-channel read successfully | 15 sec | SPI_Service
TC-MAX-LIB-012 | MAX30009 Magnitude Calculation | Calculate impedance magnitude from I/Q | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | For I=100, Q=100: magnitude = 141.42 | Magnitude calculated correctly | 10 sec | SPI_Service
TC-MAX-LIB-013 | MAX30009 Phase Calculation | Calculate impedance phase from I/Q | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | For I=100, Q=100: phase = 45Â° | Phase calculated correctly | 10 sec | SPI_Service
TC-MAX-LIB-014 | MAX30009 Load Calibration File | Load calibration from JSON file | MAX30009 Library | CM4 Linux | pytest | No | HIGH | calib/0_0.json parsed, data loaded | Calibration loaded successfully | 30 sec | JSON_Protocol + SPI_Service
TC-MAX-LIB-015 | MAX30009 Apply Calibration | Apply calibration correction to measurement | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | Raw=100, gain=1.05 â†’ calibrated=105 | Calibration applied correctly | 10 sec | SPI_Service
TC-MAX-LIB-016 | MAX30009 FIFO Count Read | Read number of samples in FIFO | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | FIFO status register read, count returned | FIFO count read correctly | 10 sec | SPI_Service
TC-MAX-LIB-017 | MAX30009 FIFO Read | Read multiple samples from FIFO | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | 10 I/Q pairs (20 values) read via burst read | FIFO read successful | 20 sec | SPI_Service
TC-MAX-LIB-018 | MAX30009 FIFO Clear | Clear FIFO buffer | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | FIFO reset register written, count = 0 after clear | FIFO cleared successfully | 10 sec | SPI_Service
TC-MAX-LIB-019 | MAX30009 Status Register Read | Read device status register | MAX30009 Library | Laptop/CM4 | gtest | No | MEDIUM | Status register (0x00) read, all bits returned | Status read successfully | 10 sec | SPI_Service
TC-MAX-LIB-020 | MAX30009 Error Detection | Detect ADC errors (saturation , PLL unlock) | MAX30009 Library | Laptop/CM4 | gtest | No | HIGH | Error bits detected in status register | Errors detected correctly | Errors detected correctly | SPI_Service
TC-MAX-PROC-001 | MAX30009 Process Constructor | Create MAX30009_process object | MAX30009 Process | Laptop/CM4 | gtest | No | CRITICAL | Object created, FIFO allocated, streaming=false | Process object created | 10 sec | SPI_Service
TC-MAX-PROC-002 | MAX30009 Process Initialization | Initialize MAX30009 process | MAX30009 Process | Laptop/CM4 | gtest | No | CRITICAL | I2C opened, registers configured, FIFO ready | Process initialized | 20 sec | SPI_Service
TC-MAX-PROC-003 | MAX30009 Configure via JSON | Configure MAX30009 via JSON command | MAX30009 Process | CM4 Linux | pytest | No | HIGH | JSON config applied, device configured, success returned | Configuration successful | 30 sec | JSON_Protocol + SPI_Service
TC-MAX-PROC-004 | MAX30009 Get Configuration | Retrieve current configuration as JSON | MAX30009 Process | CM4 Linux | pytest | No | MEDIUM | JSON with current settings returned | Configuration retrieved | 20 sec | JSON_Protocol + SPI_Service
TC-MAX-PROC-005 | MAX30009 Start Streaming | Start impedance data streaming | MAX30009 Process | CM4 Linux | pytest | No | CRITICAL | Acquisition thread started, is_streaming()=true | Streaming started | 30 sec | SPI_Service
TC-MAX-PROC-006 | MAX30009 Stop Streaming | Stop impedance data streaming | MAX30009 Process | CM4 Linux | pytest | No | CRITICAL | Thread stopped cleanly, is_streaming()=false | Streaming stopped | 30 sec | SPI_Service
TC-MAX-PROC-007 | MAX30009 Circular Buffer FIFO | Test internal circular buffer operations | MAX30009 Process | Laptop/CM4 | gtest | No | HIGH | 100 samples written, 50 read, FIFO pointers correct | FIFO operations correct | 30 sec | SPI_Service
TC-MAX-PROC-008 | MAX30009 Read Samples | Read samples from process FIFO | MAX30009 Process | CM4 Linux | pytest | No | HIGH | read_samples(10) returns 10 samples, FIFO updated | Samples read successfully | 20 sec | SPI_Service
TC-MAX-PROC-009 | MAX30009 Load Calibration Set | Load all 85 calibration files | MAX30009 Process | CM4 Linux | pytest | No | MEDIUM | All 85 files loaded into calibration table | Calibration set loaded | 2 min | SPI_Service
TC-MAX-PROC-010 | MAX30009 Select Calibration | Select specific calibration from table | MAX30009 Process | Laptop/CM4 | gtest | No | MEDIUM | Calibration for current=2, freq=5 selected from table | Correct calibration selected | 15 sec | SPI_Service
TC-MAX-PROC-011 | MAX30009 External MUX Control | Control external multiplexer for channel selection | MAX30009 Process | CM4 Linux | pytest | Yes - PCB | MEDIUM | GPIO pins set for channel 3 (binary 011) | MUX channel selected | 30 sec | SPI_Service
TC-MAX-PROC-012 | MAX30009 Status Query | Query process status via JSON | MAX30009 Process | CM4 Linux | pytest | No | MEDIUM | JSON with streaming state, sample_count, errors | Status returned correctly | 20 sec | JSON_Protocol + SPI_Service
TC-MAX-PROC-013 | MAX30009 I2C Error Handling | Handle I2C timeout gracefully | MAX30009 Process | Laptop/CM4 | gtest | No | HIGH | Error logged, get_last_error() returns description | I2C error handled | 20 sec | SPI_Service
TC-MAX-PROC-014 | MAX30009 Thread Safety | Verify concurrent FIFO access is thread-safe | MAX30009 Process | Laptop/CM4 | gtest | No | MEDIUM | No data corruption during concurrent read/write | Thread-safe operation | 1 min | SPI_Service
TC-MAX-MUX-001 | External MUX GPIO Init | Initialize GPIO pins for external MUX | MAX30009 MUX | Laptop/CM4 | gtest | No | MEDIUM | 4 GPIO pins (A0, A1, A2, EN) configured as outputs | MUX GPIOs initialized | 15 sec | SPI_Service
TC-MAX-MUX-002 | External MUX Channel Selection | Select MUX channel via GPIO | MAX30009 MUX | Laptop/CM4 | gtest | No | MEDIUM | Channel 5 (101 binary): A0=HIGH, A1=LOW, A2=HIGH | Channel selected correctly | 15 sec | SPI_Service
TC-MAX-MUX-003 | External MUX Enable Disable | Enable/disable external MUX | MAX30009 MUX | Laptop/CM4 | gtest | No | MEDIUM | EN pin toggles correctly for enable/disable | MUX enabled/disabled | 15 sec | SPI_Service
TC-MAX-MUX-004 | External MUX Invalid Channel | Handle invalid channel selection | MAX30009 MUX | Laptop/CM4 | gtest | No | LOW | Channel 8 (out of range 0-7) rejected or clamped | Invalid channel handled | 10 sec | SPI_Service
TC-ADS-LIB-001 | ADS1293 SPI Initialization | Initialize SPI communication with ADS1293 | ADS1293 Library | Laptop/CM4 | gtest | No | CRITICAL | SPI /dev/spidev0.0 opened, mode MODE_1, speed 5MHz | SPI initialized successfully | 15 sec | SPI_Service
TC-ADS-LIB-002 | ADS1293 Hardware Reset | Perform hardware reset via GPIO | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | Reset GPIO toggled LOWâ†’HIGH with â‰¥10ms delay | Hardware reset successful | 20 sec | SPI_Service
TC-ADS-LIB-003 | ADS1293 Software Reset | Perform software reset via register | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | RESET command written, registers return to defaults | Software reset successful | 15 sec | SPI_Service
TC-ADS-LIB-004 | ADS1293 Register Read | Read register via SPI | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | SPI read transaction correct (opcode + address) | Register read successful | 15 sec | SPI_Service
TC-ADS-LIB-005 | ADS1293 Register Write | Write to register via SPI | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | SPI write transaction correct, readback verifies | Register write successful | 15 sec | SPI_Service
TC-ADS-LIB-006 | ADS1293 Burst Register Read | Read multiple registers in one transaction | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | Burst read of 6 bytes (all channel data) successful | Burst read successful | 20 sec | SPI_Service
TC-ADS-LIB-007 | ADS1293 Read Device ID Hardware | Read device ID from real hardware | ADS1293 Library | CM4 Linux | pytest | Yes - PCB | CRITICAL | Device ID = 0xAD (or per datasheet) | Correct device ID | 1 min | SPI_Service
TC-ADS-LIB-008 | ADS1293 Configure Channels | Enable/disable ECG channels | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | Channel enable bits (CHxEN) set correctly for 3 channels | Channels configured | 15 sec | SPI_Service
TC-ADS-LIB-009 | ADS1293 Set Sampling Rate | Configure sampling rate (125-1000Hz) | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | DR[2:0] bits set correctly for 125, 250, 500, 1000 Hz | Sampling rate set | 20 sec | JSON_Protocol + SPI_Service
TC-ADS-LIB-010 | ADS1293 Set Gain | Configure ADC gain (6,9,12db) | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | GAIN[1:0] bits set correctly |  | GAIN[1:0] bits set correctly | JSON_Protocol + SPI_Service
TC-ADS-LIB-011 | ADS1293 Enable Leads-Off Detection | Enable leads-off detection feature | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | LOD_EN bit set in config register | Leads-off enabled | 15 sec | SPI_Service
TC-ADS-LIB-012 | ADS1293 Enable RLD | Enable Right Leg Drive (RLD) circuit | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | RLD_EN bit set in config register | RLD enabled | 15 sec | SPI_Service
TC-ADS-LIB-013 | ADS1293 Start Conversion | Start ADC conversion | ADS1293 Library | Laptop/CM4 | gtest | No | CRITICAL | START bit set, ADC begins converting | Conversion started | 15 sec | SPI_Service
TC-ADS-LIB-014 | ADS1293 Stop Conversion | Stop ADC conversion | ADS1293 Library | Laptop/CM4 | gtest | No | CRITICAL | START bit cleared, conversion stopped | Conversion stopped | 15 sec | SPI_Service
TC-ADS-LIB-015 | ADS1293 Data Ready Check | Check DRDY status for new data | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | DRDY pin or status bit read correctly | Data ready check functional | 15 sec | SPI_Service
TC-ADS-LIB-016 | ADS1293 Read Single Channel | Read data from one ECG channel | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | 3 bytes (24-bit) read, sign-extended to int32_t | Channel data read | 20 sec | SPI_Service
TC-ADS-LIB-017 | ADS1293 Read All Channels | Read data from all 3 ECG channels | ADS1293 Library | Laptop/CM4 | gtest | No | HIGH | All 3 channels read in one call | All channel data read | 25 sec | SPI_Service
TC-ADS-LIB-018 | ADS1293 Status Register Read | Read status register | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | Status register read, all bits returned | Status read successfully | 15 sec | SPI_Service
TC-ADS-LIB-019 | ADS1293 Leads-Off Detection Check | Check leads-off status | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | LOD status bits read, electrode status returned | Leads-off check functional | 15 sec | SPI_Service
TC-ADS-LIB-020 | ADS1293 Fault Detection | Detect ADC fault conditions | ADS1293 Library | Laptop/CM4 | gtest | No | MEDIUM | FAULT bit read, saturation/error detected | Fault detection functional | 15 sec | SPI_Service
TC-ADS-LIB-021 | ADS1293 Enter Standby Mode | Put device in low-power standby | ADS1293 Library | Laptop/CM4 | gtest | No | LOW | STANDBY command written, power reduced | Standby mode entered | 15 sec | SPI_Service
TC-ADS-LIB-022 | ADS1293 Wake Up from Standby | Wake device from standby mode | ADS1293 Library | Laptop/CM4 | gtest | No | LOW | WAKEUP command written, device operational | Wake up successful | 15 sec | SPI_Service
TC-ADS-PROC-001 | ADS1293 Process Constructor | Create ADS1293_process object | ADS1293 Process | Laptop/CM4 | gtest | No | CRITICAL | FIFOs allocated (size=3000), acquiring=false | Process object created | 10 sec | SPI_Service
TC-ADS-PROC-002 | ADS1293 Process Initialization | Initialize ADS1293 process | ADS1293 Process | Laptop/CM4 | gtest | No | CRITICAL | SPI opened, GPIOs configured, device ID read | Process initialized | 20 sec | SPI_Service
TC-ADS-PROC-003 | ADS1293 Configure via JSON | Configure ADS1293 via JSON command | ADS1293 Process | CM4 Linux | pytest | No | HIGH | JSON config applied: rate=500, channels=3, gain=12 | Configuration successful | 30 sec | JSON_Protocol + SPI_Service
TC-ADS-PROC-004 | ADS1293 Start Acquisition | Start ECG data acquisition | ADS1293 Process | CM4 Linux | pytest | No | CRITICAL | Acquisition thread started, is_acquiring()=true | Acquisition started | 30 sec | SPI_Service
TC-ADS-PROC-005 | ADS1293 Stop Acquisition | Stop ECG data acquisition | ADS1293 Process | CM4 Linux | pytest | No | CRITICAL | Thread stopped cleanly, is_acquiring()=false | Acquisition stopped | 30 sec | SPI_Service
TC-ADS-PROC-006 | ADS1293 FIFO Write | Write samples to internal FIFO | ADS1293 Process | Laptop/CM4 | gtest | No | HIGH | Sample written to all 3 channel FIFOs, write_pos increments | FIFO write successful | 20 sec | SPI_Service
TC-ADS-PROC-007 | ADS1293 FIFO Read | Read samples from internal FIFO | ADS1293 Process | Laptop/CM4 | gtest | No | HIGH | 10 samples read returns 30 values (3 ch Ã— 10), pointers updated | FIFO read successful | 25 sec | SPI_Service
TC-ADS-PROC-008 | ADS1293 FIFO Wraparound | Test circular buffer wraparound | ADS1293 Process | Laptop/CM4 | gtest | No | MEDIUM | Write 3100 samples (size=3000), write_pos wraps to 100 | Wraparound successful | 30 sec | SPI_Service
TC-ADS-PROC-009 | ADS1293 FIFO Full Condition | Detect FIFO full condition | ADS1293 Process | Laptop/CM4 | gtest | No | MEDIUM | FIFO full at 3000 samples, overflow handled | Full condition detected | 20 sec | SPI_Service
TC-ADS-PROC-010 | ADS1293 FIFO Empty Condition | Detect FIFO empty condition | ADS1293 Process | Laptop/CM4 | gtest | No | MEDIUM | Read from empty FIFO returns empty vector | Empty condition detected | 15 sec | SPI_Service
TC-ADS-PROC-011 | ADS1293 Sync Mark Insertion | Insert sync mark into data stream | ADS1293 Process | Laptop/CM4 | gtest | No | HIGH | SYNC_MARK_MAGIC_NUM (-99999) inserted in all 3 FIFOs | Sync mark inserted | 20 sec | SPI_Service
TC-ADS-PROC-012 | ADS1293 Get Available Samples | Query number of available samples | ADS1293 Process | Laptop/CM4 | gtest | No | MEDIUM | Write 50, read 30 â†’ available = 20 | Available samples correct | 15 sec | SPI_Service
TC-ADS-PROC-013 | ADS1293 Acquisition Loop | Test acquisition thread operation | ADS1293 Process | CM4 Linux | pytest | No | HIGH | Samples accumulate in FIFO during 1-second run | Acquisition loop functional | 1 min | SPI_Service
TC-ADS-PROC-014 | ADS1293 Thread Safety | Verify FIFO thread-safety | ADS1293 Process | Laptop/CM4 | gtest | No | MEDIUM | Concurrent read/write doesn't corrupt data, mutex works | Thread-safe operation | 1 min | SPI_Service
TC-ADS-PROC-015 | ADS1293 Get Status | Query process status via JSON | ADS1293 Process | CM4 Linux | pytest | No | MEDIUM | JSON with acquiring state, sample_count, FIFO usage | Status returned correctly | 20 sec | JSON_Protocol + SPI_Service
TC-ADS-PROC-016 | ADS1293 SPI Error Handling | Handle SPI timeout gracefully | ADS1293 Process | Laptop/CM4 | gtest | No | HIGH | Error logged, get_last_error() returns description | SPI error handled | 20 sec | SPI_Service
TC-WS-DRV-001 | WS2812 Structure Initialization | Initialize ws2811_t structure | WS2812 Driver | Laptop/CM4 | gtest | No | MEDIUM | freq=800000, count=30, brightness=255 | Structure initialized | 10 sec | SPI_Service
TC-WS-DRV-002 | WS2812 Driver Init | Initialize WS2812 driver | WS2812 Driver | Laptop/CM4 | gtest | No | HIGH | DMA allocated, PWM configured, returns WS2811_SUCCESS | Driver initialized | 20 sec | SPI_Service
TC-WS-DRV-003 | WS2812 Driver Fini | Clean up WS2812 driver | WS2812 Driver | Laptop/CM4 | gtest | No | HIGH | DMA freed, PWM disabled, no memory leaks | Driver cleaned up | 15 sec | SPI_Service
TC-WS-DRV-004 | WS2812 Set LED Color | Set single LED to specific color | WS2812 Driver | Laptop/CM4 | gtest | No | HIGH | LED 0 set to 0xFF0000 (red) | LED color set | 10 sec | SPI_Service
TC-WS-DRV-005 | WS2812 Get LED Color | Read current LED color | WS2812 Driver | Laptop/CM4 | gtest | No | MEDIUM | Get LED 0 returns 0xFF0000 | LED color retrieved | 10 sec | SPI_Service
TC-WS-DRV-006 | WS2812 Render | Render LED data to strip | WS2812 Driver | Laptop/CM4 | gtest | No | CRITICAL | DMA triggered, PWM generates signal | Render successful | 20 sec | SPI_Service
TC-WS-DRV-007 | WS2812 Timing Validation | Validate timing with logic analyzer | WS2812 Driver | CM4 Linux | pytest | Yes - Analyzer | LOW | T0H=350ns, T1H=700ns, reset=50Âµs (Â±150ns) | Timing within spec | 5 min | SPI_Service
TC-WS-DRV-008 | WS2812 Multiple LEDs | Set 30 LEDs to different colors | WS2812 Driver | CM4 Linux | pytest | Yes - PCB | MEDIUM | All 30 LEDs change colors (visual) | All LEDs functional | 1 min | SPI_Service
TC-WS-DRV-009 | WS2812 Brightness Control | Control global brightness | WS2812 Driver | Laptop/CM4 | gtest | No | MEDIUM | Brightness=128 (50%), LED 0xFFâ†’0x7F | Brightness control works | 15 sec | SPI_Service
TC-WS-DRV-010 | WS2812 Out of Bounds Index | Handle invalid LED index | WS2812 Driver | Laptop/CM4 | gtest | No | LOW | LED index 50 (count=30) rejected or clamped | Bounds checking works | 10 sec | SPI_Service
TC-WS-PROC-001 | WS2812 Process Constructor | Create WS2812_process object | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | ws2811_t initialized with count=30, GPIO=18 | Process object created | 10 sec | SPI_Service
TC-WS-PROC-002 | WS2812 Process Initialization | Initialize WS2812 process | WS2812 Process | Laptop/CM4 | gtest | No | HIGH | ws2811_init() called successfully | Process initialized | 20 sec | SPI_Service
TC-WS-PROC-003 | WS2812 Set LED Count | Change LED strip length | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | LED count changed to 60, buffers reallocated | LED count updated | 15 sec | SPI_Service
TC-WS-PROC-004 | WS2812 Set Brightness | Set global brightness level | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | Brightness set to 200 | Brightness updated | 10 sec | SPI_Service
TC-WS-PROC-005 | WS2812 Set Single LED Color | Set color of one LED | WS2812 Process | Laptop/CM4 | gtest | No | HIGH | LED 0 set to red (255,0,0) | Single LED color set | 15 sec | SPI_Service
TC-WS-PROC-006 | WS2812 Set All LEDs Color | Set all LEDs to same color | WS2812 Process | Laptop/CM4 | gtest | No | HIGH | All 30 LEDs set to green (0,255,0) | All LEDs color set | 15 sec | SPI_Service
TC-WS-PROC-007 | WS2812 Clear All LEDs | Turn off all LEDs | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | All LEDs set to 0x000000 (off) | All LEDs cleared | 15 sec | SPI_Service
TC-WS-PROC-008 | WS2812 Solid Color Pattern | Display solid color pattern | WS2812 Process | CM4 Linux | pytest | Yes - PCB | HIGH | All LEDs turn red (visual verification) | Solid red displayed | 30 sec | SPI_Service
TC-WS-PROC-009 | WS2812 Blink Pattern | Display blink pattern | WS2812 Process | CM4 Linux | pytest | Yes - PCB | MEDIUM | LEDs blink blue on/off every 500ms (visual) | Blink pattern functional | 1 min | SPI_Service
TC-WS-PROC-010 | WS2812 Fade Pattern | Display fade transition pattern | WS2812 Process | CM4 Linux | pytest | Yes - PCB | MEDIUM | Smooth redâ†’blue fade over 2s (visual) | Fade pattern functional | 2 min | SPI_Service
TC-WS-PROC-011 | WS2812 Rainbow Cycle Pattern | Display rainbow cycle pattern | WS2812 Process | CM4 Linux | pytest | Yes - PCB | MEDIUM | LEDs cycle through rainbow colors (visual) | Rainbow pattern functional | 1 min | SPI_Service
TC-WS-PROC-012 | WS2812 Configure via JSON | Configure LED via JSON command | WS2812 Process | CM4 Linux | pytest | No | HIGH | JSON config applied: solid red pattern | JSON configuration works | 30 sec | JSON_Protocol + SPI_Service
TC-WS-PROC-013 | WS2812 Get Status | Query LED status via JSON | WS2812 Process | CM4 Linux | pytest | No | MEDIUM | JSON with LED count, pattern, colors | Status returned correctly | 20 sec | JSON_Protocol + SPI_Service
TC-WS-PROC-014 | WS2812 Pattern Thread | Verify pattern thread spawns correctly | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | Blink pattern thread spawned, running_pattern=true | Pattern thread functional | 20 sec | SPI_Service
TC-WS-PROC-015 | WS2812 Stop Pattern | Stop running pattern | WS2812 Process | Laptop/CM4 | gtest | No | MEDIUM | Pattern thread stopped, new pattern applied | Pattern stopped correctly | 20 sec | SPI_Service
TC-WS-PROC-016 | WS2812 Show Function | Update LEDs with current data | WS2812 Process | CM4 Linux | pytest | Yes - PCB | HIGH | ws2811_render() called, LEDs update | LEDs updated visually | 30 sec | SPI_Service
TC-PWR-PROC-001 | Power Process Constructor | Create PWRCNTR_process object | Power Control | Laptop/CM4 | gtest | No | MEDIUM | GPIO drivers initialized, SMBus created | Process object created | 10 sec | Power_Service
TC-PWR-PROC-002 | Power Process Initialization | Initialize power control process | Power Control | Laptop/CM4 | gtest | No | HIGH | GPIOs exported, SMBus opened to 0x0B | Process initialized | 20 sec | Power_Service
TC-PWR-PROC-003 | Set Power Rail Enable | Enable power rail via GPIO | Power Control | Laptop/CM4 | gtest | No | HIGH | GPIO set HIGH for ADS1293 rail | Power rail enabled | 15 sec | Power_Service
TC-PWR-PROC-004 | Set Power Rail Disable | Disable power rail via GPIO | Power Control | Laptop/CM4 | gtest | No | HIGH | GPIO set LOW for MAX30009 rail | Power rail disabled | 15 sec | Power_Service
TC-PWR-PROC-005 | Get Power Rail Status | Query power rail status | Power Control | Laptop/CM4 | gtest | No | MEDIUM | Status query returns true for enabled rail | Status query works | 10 sec | Power_Service
TC-PWR-PROC-006 | Read Battery Voltage | Read battery voltage via SMBus | Power Control | Laptop/CM4 | gtest | No | HIGH | Voltage 3.0-4.2V read from SMBus | Voltage read successfully | 20 sec | Power_Service
TC-PWR-PROC-007 | Read Battery Current | Read battery current via SMBus | Power Control | Laptop/CM4 | gtest | No | MEDIUM | Current 0-2000mA read from SMBus | Current read successfully | 20 sec | Power_Service
TC-PWR-PROC-008 | Read Battery SoC | Read battery state of charge | Power Control | Laptop/CM4 | gtest | No | MEDIUM | SoC 0-100% read from SMBus | SoC read successfully | 20 sec | Power_Service
TC-PWR-PROC-009 | Read Battery Capacity | Read battery capacity in mAh | Power Control | Laptop/CM4 | gtest | No | LOW | Capacity in mAh read from SMBus | Capacity read successfully | 20 sec | Power_Service
TC-PWR-PROC-010 | Configure via JSON | Configure power control via JSON | Power Control | CM4 Linux | pytest | No | MEDIUM | JSON config applied to power rails | Configuration successful | 30 sec | JSON_Protocol + Power_Service
TC-PWR-PROC-011 | Get Status | Query power control status | Power Control | CM4 Linux | pytest | No | MEDIUM | JSON with all rail states returned | Status returned correctly | 20 sec | Power_Service
TC-PWR-PROC-012 | Query Battery | Query complete battery telemetry | Power Control | CM4 Linux | pytest | No | HIGH | JSON with voltage, current, SoC, capacity | Battery query successful | 30 sec | Power_Service
TC-PWR-PROC-013 | Power Cycle Sequence | Execute power disableâ†’waitâ†’enable sequence | Power Control | CM4 Linux | pytest | Yes - PCB | HIGH | Rail disabled, 2s wait, rail enabled, device resets | Power cycle successful | 3 min | Power_Service
TC-PWR-PROC-014 | SMBus Timeout Error Handling | Handle SMBus communication timeout | Power Control | Laptop/CM4 | gtest | No | MEDIUM | Error logged, default value returned | SMBus timeout handled | 20 sec | Power_Service
TC-SMB-DRV-001 | SMBus Driver Open | Open SMBus device | SMBus Driver | Laptop/CM4 | gtest | No | HIGH | I2C device /dev/i2c-1 opened at address 0x0B | SMBus opened successfully | 15 sec | Power_Service
TC-SMB-DRV-002 | SMBus Read Byte | Read byte from SMBus register | SMBus Driver | Laptop/CM4 | gtest | No | HIGH | Correct SMBus read transaction | Byte read successfully | 15 sec | Power_Service
TC-SMB-DRV-003 | SMBus Read Word | Read word (16-bit) from SMBus | SMBus Driver | Laptop/CM4 | gtest | No | MEDIUM | 2 bytes read, little-endian conversion | Word read successfully | 15 sec | Power_Service
TC-SMB-DRV-004 | SMBus Write Byte | Write byte to SMBus register | SMBus Driver | Laptop/CM4 | gtest | No | MEDIUM | SMBus write command sent | Byte written successfully | 15 sec | Power_Service
TC-SMB-DRV-005 | SMBus Real Battery Communication | Communicate with real battery IC | SMBus Driver | CM4 Linux | pytest | Yes - PCB | HIGH | Voltage/current read from real battery IC, values plausible | Battery communication works | 2 min | Power_Service